Section III: Returnable Bidding Forms

eSourcing reference:  RFQ/2026/62289

Note to Bidders: The following returnable forms are part of this RFQ and must be completed and returned by bidders as part of their quotation. Instructions to complete each Form are highlighted in blue in each Form. Please complete the Returnable Bidding Forms as instructed and return them as part of your quotation by uploading them against their specific Document Checklist in the UNOPS eSourcing system.

This Section comprises the following Returnable Bidding Forms:

- Form A: Quotation Submission Form

- Form B: Price Schedule Form

- Form C: Technical Quotation Form

- Form D: Previous Experience Form

- Form E: Self Disclosure Form

- Form F: Joint Venture Partner Information Form

- Form G: Manufacturer’s authorization form

Form A: Quotation submission form

Bidders are requested to complete this form, sign it and return it as part of their bid submission. The bidder shall fill in this form in accordance with the instructions indicated. No alterations to its format shall be permitted and no substitutions shall be accepted.

Date: [Insert submission date]

Subject: Quotation for the supply of [Insert a brief description of goods/services] in [Name of country/city], RFQ Case No. [Insert RFQ ref number], dated [insert date]

We, the undersigned, declare that:

  - We offer to supply in conformity with the bidding documents, including the UNOPS General Conditions of Contract;

  - Our quotation shall be valid for the period of time of [insert number of days which shall not be less than the specified in the Tender Particulars section, Period of Validity of Quotations] from the date fixed for the submission deadline as set out in the RFQ, and it shall remain binding upon us and may be accepted at any time before the expiration of that period;

  - We have no conflict of interest in any activity that would put it, if selected for this assignment, in a conflict of interest with UNOPS [If you have any actual or potential conflict of interest as defined in Article 3 of Section I: Instructions to Bidders, please disclose it here];;

  - Our firm confirms that the offeror and sub-contractors have not been associated, or had been involved in any way, directly or indirectly, with the preparation of the design, terms of references and/or other documents used as a part of this solicitation;

  - Our firm, its affiliates or subsidiaries—including any subcontractors or suppliers for any part of the Contract—has not been declared ineligible by UNOPS, nor is included in the suspended/ineligibility list of the UN/PD, other UN Agencies, the UN Security Council, and the World Bank, in accordance with Instructions to Bidders Article 3, Eligibility;

  - We embrace the UN Supplier Code of Conduct and adhere to the principles of the UN Global Compact;

  - We have not declared bankruptcy, are not involved in bankruptcy or receivership proceedings, and there is no judgment or pending legal action against them that could impair their operations in the foreseeable future;

  - We have not offered and will not offer fees, gifts and/or favours of kind in exchange for this RFQ and will not engage in any such activity during the performance of any Contract awarded.

I, the undersigned, certify that I am duly authorized by [insert full name of bidder] to sign this quotation and bind [insert full name of bidder] should UNOPS accept this quotation:

Name: [complete]

Title: [complete]

Date: [complete]

Signature: _____________________________________________________________

Provide the name and contact information for the primary contact from your company for this quotation:

Name: [complete]

Title: [complete]

Email address: [complete]

Telephone: [complete]

# Form B: Price Schedule Form

Bidders shall fill in this Price Schedule Form in accordance with the instructions indicated.

**Please fill in all your prices in the attached excel sheet and send it as (excel book) along with the below table**

Pricing Methodology:

- In addition to DAP price bidder to add a breakdown of prices: cost of goods cleared for export (FCA) separate from transport and insurance.

- The Supplier shall, within fourteen (14) days of Contract signature, provide a contingency plan for alternative routing (e.g., land corridors or bypass ports) should the primary maritime route be restricted.

- To confirm whether the primary components and finished goods are sourced from regions not currently subject to maritime blockades or significant transit disruptions in the Strait of Hormuz or Red Sea (Geographic Origin Declaration).

RFQ reference no: [insert RFQ reference No.]

| Currency | USD |
| --- | --- |

Bills of quantity Lot# 1: Supply & installation and commissioning of general laboratory medical equipment

| Item No | Description |  | Qty | Locations and Distributions | Unit Price DAP | Total Price DAP |
| --- | --- | --- | --- | --- | --- | --- |
|  | Biological safety Cabinet class II, Type A2 |  | 2 | National Central Lab, Lahj | insert | insert |
|  | Real Time PCR Machine |  | 1 | National Central Lab, Lahj | insert | insert |
|  | Analytical Balance with Air Shield Cabinet |  | 1 | National Central Lab, Lahj | insert | insert |
|  | Hot Air Oven 50 L |  | 2 | National Central Lab, Lahj: 1 Dr. Ameen Nasher, Aden Building A:1 | insert | insert |
|  | Shaker |  | 3 | Ataq TFC, Shabwah: 1 Al-Mansoura TFC, Aden:1 Dr. Ameen Nasher, Aden Building A:1 | insert | insert |
|  | Mixer, vortex |  | 4 | Ataq TFC, Shabwah: 1 Al-Mansoura TFC, Aden:1 Hais TFC, Al Hudayda: 1 Dr. Ameen Nasher Building A, Aden:1 | insert | insert |
|  | Elisa Reader +Washer |  | 2 | Ataq TFC, Shabwah:1 Al-Mansoura TFC, Aden:1 | insert | insert |
|  | BOD Incubator |  | 2 | Ataq TFC, Shabwah:1 Al-Mansoura TFC, Aden:1 | insert | insert |
|  | Magnetic Stirrer with Hot Plate |  | 1 | Dr. Ameen Nasher, Aden Building A:1 | insert | insert |
|  | Water-Jacketed CO2 Incubators |  | 1 | Dr. Ameen Nasher, Aden Building A:1 | insert | insert |
|  | B2 Biological Safety Cabinet, floor-standing |  | 2 | Dr. Ameen Nasher, Aden Building A:2 | insert | insert |
|  | Fume Hoods, Floor-Mounted, width 1200mm |  | 2 | Dr. Ameen Nasher, Aden Building A:2 | insert | insert |
|  | Laminar Air Flow, Class II, B2 |  | 2 | Dr. Ameen Nasher, Aden Building A:2 | insert | insert |
|  | Bunsen Burner |  | 2 | Ataq TFC: 1 Al-Mansoura TFC, Aden:1 | insert | insert |
|  | Hoods, Canopy |  | 2 | Dr. Ameen Nasher Building A:2 | insert | insert |
|  |  | Total Price |  |  |  | insert |

Payment terms 30 days accepted: ☐ Yes

Bidder’s discount for accelerated payment: ____% of total firm price for each calendar day less than thirty (30) days

List of subcontractors or suppliers

Bidder must identify the names of all subcontractors/suppliers who will be providing good/services under this Contract and the type of work being subcontracted, if applicable.

- _[Full legal name and address of subcontractors]___________

- _________________________________________________

- _________________________________________________

I, the undersigned, certify that I am duly authorized by [insert full name of Bidder] to sign this quotation and bind [insert full name of Bidder] should UNOPS accept this quotation:

Name	: _____________________________________________________________

Title	: _____________________________________________________________

Date	: _____________________________________________________________

Signature	: _____________________________________________________________

Form C: Technical Quotation Form

RFQ reference no: [insert RFQ reference No.]

Name of Bidder: [insert name of Bidder]

Bidders are required to complete the Comparative Data Tables included in Section II: Schedule of Requirements to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.

Technical specifications for goods – Comparative Data Table

| Item No | UNOPS minimum technical requirements | Qty | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete |
| --- | --- | --- | --- | --- |
| 1 | Biological safety Cabinet class II, Type A2 |  |  |  |
|  | Name of Manufacturer | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | provides personnel, environmental and product protection |  | ☐ Yes ☐ No |  |
|  | Internal working area dimensions approx. (W x D x H, in mm): 1500 x 600 x 750, permissible deviation in width +/-10%, in depth and height +/-15% |  | ☐ Yes ☐ No |  |
|  | 70% air recirculated; 30% exhausted from a common plenum to the room |  | ☐ Yes ☐ No |  |
|  | Stainless steel cabinet housing with antibacterial coating |  | ☐ Yes ☐ No |  |
|  | Front window can be opened for easy cleaning |  | ☐ Yes ☐ No |  |
|  | In-built timer: available |  | ☐ Yes ☐ No |  |
|  | Display with air flow volume and speed values, HEPA filter operating time in %, total operating time, UV and fluorescent lamps operating time values |  | ☐ Yes ☐ No |  |
|  | Audible and visual alarm system for compromised airflow |  | ☐ Yes ☐ No |  |
|  | Automatic UV lamp turn-off when the front window is opened |  | ☐ Yes ☐ No |  |
|  | Audio and visual notifications for air flow speed, system availability, front window position, alarms for lamp replacement or maintenance |  | ☐ Yes ☐ No |  |
|  | EN 12469 certificate: available |  | ☐ Yes ☐ No |  |
|  | Compatible stand for BSC installation, featuring casters and a position-locking function |  | ☐ Yes ☐ No |  |
|  | Average inflow velocity of approx. 0,5 m/s (100 fpm) or better |  | ☐ Yes ☐ No |  |
|  | Average down flow velocity approx. 0,3 m/s (60 fpm) or better |  | ☐ Yes ☐ No |  |
|  | Exhaust and intake HEPA or ULPA filters with filter efficiency of at least 99,99% MPPS |  | ☐ Yes ☐ No |  |
|  | At least 304 type stainless steel working surface |  | ☐ Yes ☐ No |  |
|  | Removable work tray (work surface) for easy cleaning and decontamination |  | ☐ Yes ☐ No |  |
|  | at least 2 waterproof EU standard sockets |  | ☐ Yes ☐ No |  |
|  | Noise level maximum 58 dB (A) in standard mode |  | ☐ Yes ☐ No |  |
|  | Lighting in working area at least 1250 lux |  | ☐ Yes ☐ No |  |
|  | Front glass thickness at least 6 mm |  | ☐ Yes ☐ No |  |
|  | Supplied with |  | ☐ Yes ☐ No |  |
|  | 1 set of spare intake and exhaust HEPA or ULPA filter(s) |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent. |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |
| 2 | Real Time PCR Machine |  |  |  |
|  | Name of Manufacturer | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Obtain Precise Quantification & Multiplex Target Discrimination |  | ☐ Yes ☐ No |  |
|  | Thermal cycling system should be Peltier-based, 96-well block |  | ☐ Yes ☐ No |  |
|  | Obtain Accurate Results with Low Volumes of Sample & Reagents, SPECIFY |  | ☐ Yes ☐ No |  |
|  | PCR tests in separate tubes, strips and plates with compatible PCR Tubes. |  | ☐ Yes ☐ No |  |
|  | Excitation beam - white LED: Available |  | ☐ Yes ☐ No |  |
|  | Excitation channels (filters): At least 6 |  | ☐ Yes ☐ No |  |
|  | Detection channels (filters): At least 6 |  | ☐ Yes ☐ No |  |
|  | Multiplex assays for at least 6 targets in a single tube |  | ☐ Yes ☐ No |  |
|  | Fully compatible with all dyes |  | ☐ Yes ☐ No |  |
|  | Excitation/detection waves range 450 nm — 700 nm |  | ☐ Yes ☐ No |  |
|  | TEMPERATURE RANGE :0 °C to 100 °C |  | ☐ Yes ☐ No |  |
|  | Temperature accuracy ± 0.25°С or less |  | ☐ Yes ☐ No |  |
|  | Thermal block temperature uniformity max. ± 0.5º С |  | ☐ Yes ☐ No |  |
|  | Maximum Heating Ramp Rate 4 °С/s |  | ☐ Yes ☐ No |  |
|  | Temperature differential Range:1-30 °C |  | ☐ Yes ☐ No |  |
|  | Availability of 6 independent thermally-regulated blocks for PCR optimisation and concurrent independent tests |  | ☐ Yes ☐ No |  |
|  | Applicable for pathogen detection |  | ☐ Yes ☐ No |  |
|  | Sensitivity: Detection of 1 copy |  | ☐ Yes ☐ No |  |
|  | Colour touch screen 8" |  | ☐ Yes ☐ No |  |
|  | Ethernet, USB and WiFi ports: included |  | ☐ Yes ☐ No |  |
|  | Create, review, edit amplification program, monitor thermal cycling process and its results in real time, results analysis: Available |  | ☐ Yes ☐ No |  |
|  | Real-Time PCR Analysis software: included |  | ☐ Yes ☐ No |  |
|  | Computerized analysis of amplification results : included |  | ☐ Yes ☐ No |  |
|  | Ability to design new protocols |  | ☐ Yes ☐ No |  |
|  | No. of tests per Reagent kit should be specified |  | ☐ Yes ☐ No |  |
|  | Temperature to store the reagents must be specified. |  | ☐ Yes ☐ No |  |
|  | Stability period of the opened reagent kit should be specified |  | ☐ Yes ☐ No |  |
|  | Max . No. of tests per hour should be specified |  | ☐ Yes ☐ No |  |
|  | Cost of reagents kit should be quoted for comparative evaluation. |  | ☐ Yes ☐ No |  |
|  | All reagents must be available at any time |  | ☐ Yes ☐ No |  |
|  | Supplied with |  | ☐ Yes ☐ No |  |
|  | All components, consumables and accessories necessary for having a “ready to start” device and for performing the required quality control (QC) and calibration tests. |  | ☐ Yes ☐ No |  |
|  | Laptop with the appropriate licensed operating system, should specify the manufacturer, model ,origin and the specifications details |  | ☐ Yes ☐ No |  |
|  | Reagents for one month for 20 samples/day or as per requirement should be provided along with the machine. |  | ☐ Yes ☐ No |  |
|  | Online UPS, should specify the manufacturer ,model ,origin and the specifications details |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | "A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | Compliant with IEC 60601 or equivalent" |  | ☐ Yes ☐ No |  |
| 3 | Analytical Balance with Air Shield Cabinet |  |  |  |
|  | Name of Manufacturer | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Bench top loading |  | ☐ Yes ☐ No |  |
|  | Five segment bright LCD display with back light |  | ☐ Yes ☐ No |  |
|  | Built-in auto-calibration and testing feature. |  | ☐ Yes ☐ No |  |
|  | Built-in shock and overload protection. |  | ☐ Yes ☐ No |  |
|  | Robust metal construction and chemical resistance. |  | ☐ Yes ☐ No |  |
|  | Weighing capacity: approximately 0 - 2500 gm. |  | ☐ Yes ☐ No |  |
|  | Readability: ± 0.1g. |  | ☐ Yes ☐ No |  |
|  | Zero Adjustment |  | ☐ Yes ☐ No |  |
|  | Auto environment adjustment setting. |  | ☐ Yes ☐ No |  |
|  | The balance shall have an all-glass and easily cleaned draught shield |  | ☐ Yes ☐ No |  |
|  | Stabilization time (response time) should never exceed 1.5 sec. |  | ☐ Yes ☐ No |  |
|  | Weighing pan dimensions Approximately (WxD) 170x170 mm. |  | ☐ Yes ☐ No |  |
|  | Stainless Steel housing |  | ☐ Yes ☐ No |  |
|  | To be supplied with Dust cover and a set of calibrating weights. |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  | ☐ Yes ☐ No |  |
| 4 | Hot Air Oven |  |  |  |
|  | Name of Manufacturer | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | A tabletop, ideal for general laboratory drying applications. |  | ☐ Yes ☐ No |  |
|  | A variable microprocessor based digital temperature controller from 50 to 250 C indicatively. |  | ☐ Yes ☐ No |  |
|  | Accuracy ⁰C:± 0.1⁰C |  | ☐ Yes ☐ No |  |
|  | Double walled chamber with powder coated outer cabinet and inner chamber made of Corrosion-resistant, easy to clean, stainless-steel interior. |  | ☐ Yes ☐ No |  |
|  | Exterior remains at safe temperature (not overheated). |  | ☐ Yes ☐ No |  |
|  | Digital display with digital timer and thermometer should be provided separately. |  | ☐ Yes ☐ No |  |
|  | Timer: Adjustable |  | ☐ Yes ☐ No |  |
|  | Uniform temperature on all shelves and provided with forced air circulating noise-less fan. |  | ☐ Yes ☐ No |  |
|  | Specify Min Heating Time |  | ☐ Yes ☐ No |  |
|  | Self-diagnostic facility |  | ☐ Yes ☐ No |  |
|  | Overheat and over current protection. |  | ☐ Yes ☐ No |  |
|  | Adjustable temperature safety thermostat |  | ☐ Yes ☐ No |  |
|  | Silicon Gasket Door |  | ☐ Yes ☐ No |  |
|  | Over Temperature Audio/visual temperature alarm. |  | ☐ Yes ☐ No |  |
|  | Fast temperature recovery time. (Recovery time at 200°C never exceeds five minutes). |  | ☐ Yes ☐ No |  |
|  | Adjustable shelf heights (4 shelves to be provided) |  | ☐ Yes ☐ No |  |
|  | Temperature Monitoring |  | ☐ Yes ☐ No |  |
|  | Door features grabber-type latch. |  | ☐ Yes ☐ No |  |
|  | Minimum Capacity: never be less than 100 L |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | compliant with IEC 60601 or equivalent. |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |
| 5 | Shaker |  |  |  |
|  | Name of Manufacturer | 3 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Orbital Shaking Movement |  | ☐ Yes ☐ No |  |
|  | General Description: Clinical Lab Rotator with continuous duty motor provides smooth, quiet rotation |  | ☐ Yes ☐ No |  |
|  | Microprocessor controlled equipment. |  | ☐ Yes ☐ No |  |
|  | Maintenance-free shaking motor. |  | ☐ Yes ☐ No |  |
|  | Adjustable digital speed selection up to a maximum value not lower than 500 rpm. |  | ☐ Yes ☐ No |  |
|  | Capability to shake multi tubes 16-20 tube |  | ☐ Yes ☐ No |  |
|  | At least continuous and “touch start” operating modes. |  | ☐ Yes ☐ No |  |
|  | Capability to work with standard tubes, flasks, beakers and micro-plates found in the market. |  | ☐ Yes ☐ No |  |
|  | Maximum Load:1 Kg |  | ☐ Yes ☐ No |  |
|  | Display: Actual platform speed will be clearly |  | ☐ Yes ☐ No |  |
|  | Timer: Integral timer is set at periods between 0 – 99 min |  | ☐ Yes ☐ No |  |
|  | Alarm: Signals end of timed rotation |  | ☐ Yes ☐ No |  |
|  | Construction: Platform surface will be a non-slip material, the rim will be raised to retain slides |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | compliant with IEC 60601 or equivalent. |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |
| 6 | Vortex Mixer |  |  |  |
|  | Name of Manufacturer | 4 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Device mixing effect obtained by orbital movement |  | ☐ Yes ☐ No |  |
|  | Microprocessor controlled equipment. |  | ☐ Yes ☐ No |  |
|  | Maintenance-free shaking motor. |  | ☐ Yes ☐ No |  |
|  | Adjustable speed up to a maximum value not lower than 2200 rpm. |  | ☐ Yes ☐ No |  |
|  | Capability to shake tubes with diameters up to at least 30 mm. |  | ☐ Yes ☐ No |  |
|  | At least continuous and “touch start” operating modes. |  | ☐ Yes ☐ No |  |
|  | Capability to work with standard tubes, flasks, beakers and micro-plates found in the market. |  | ☐ Yes ☐ No |  |
|  | Supplied with |  | ☐ Yes ☐ No |  |
|  | All components and accessories necessary for having a “ready to start” |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | compliant with IEC 60601 or equivalent. |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |
| 7 | ELISA Microplate Washer and Reader |  |  |  |
|  | Name of Manufacturer | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Microplate Washer for ELISA |  | ☐ Yes ☐ No |  |
|  | RELIABALE microplate ELISA washer |  | ☐ Yes ☐ No |  |
|  | Open System |  | ☐ Yes ☐ No |  |
|  | Table Top |  | ☐ Yes ☐ No |  |
|  | Well Plate Sizes: up to 96 |  | ☐ Yes ☐ No |  |
|  | programmable |  | ☐ Yes ☐ No |  |
|  | LCD Display |  | ☐ Yes ☐ No |  |
|  | Ultrasonic Cleaning |  | ☐ Yes ☐ No |  |
|  | Auto Washing Cells |  | ☐ Yes ☐ No |  |
|  | Cell wash heads straight or angled |  | ☐ Yes ☐ No |  |
|  | Self-Clean Program |  | ☐ Yes ☐ No |  |
|  | Liquid Level Sensors for and waste bottles |  | ☐ Yes ☐ No |  |
|  | Microplate Reader for ELISA |  | ☐ Yes ☐ No |  |
|  | Open System |  | ☐ Yes ☐ No |  |
|  | Table Top |  | ☐ Yes ☐ No |  |
|  | Well Plate Sizes: up to 96 |  | ☐ Yes ☐ No |  |
|  | Temperature Incubation:37 c degree |  | ☐ Yes ☐ No |  |
|  | MEASURMENT PRINCIPLE:MONO & BIOCHROMIC |  | ☐ Yes ☐ No |  |
|  | BULB HALOGEN OR EQUIVELANT |  | ☐ Yes ☐ No |  |
|  | Programmable |  | ☐ Yes ☐ No |  |
|  | LCD Display |  | ☐ Yes ☐ No |  |
|  | Self-Diagnostics |  | ☐ Yes ☐ No |  |
|  | Needed Wavelengths: 400–630 Nm |  | ☐ Yes ☐ No |  |
|  | number of INTERFERENCE filters: 5 — 10 |  | ☐ Yes ☐ No |  |
|  | Results output (built in thermal printer) |  | ☐ Yes ☐ No |  |
|  | HL7 compatibility for HIS integration |  | ☐ Yes ☐ No |  |
|  | Electrical Characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
| 8 | BOD Incubator |  |  |  |
|  | Name of Manufacturer | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Applications being able to provide exact temperature control for microbial cultivation, biochemical assays, and environmental testing. |  | ☐ Yes ☐ No |  |
|  | Used to maintain the temperature and humidity of a sample during the biochemical oxygen demand (BOD) test. |  | ☐ Yes ☐ No |  |
|  | the BOD test measures the amount of oxygen consumed by microorganisms while decomposing organic matter in water |  | ☐ Yes ☐ No |  |
|  | The incubator must maintain a temperature between 0-60°C . |  | ☐ Yes ☐ No |  |
|  | LED display with touch keys for setting the parameters |  | ☐ Yes ☐ No |  |
|  | Temperature Range 0-60°C Temperature Accuracy ±0.5°C Temperature Fluctuation ±1°C Temperature Uniformity ±1°C Temperature Resolution ±0.5°C |  | ☐ Yes ☐ No |  |
|  | Capacity 70 L |  | ☐ Yes ☐ No |  |
|  | built-in BOD socket |  | ☐ Yes ☐ No |  |
|  | Timer Range 0-999 min |  | ☐ Yes ☐ No |  |
|  | Stainless steel inner chamber with easy to clean semicircular corners |  | ☐ Yes ☐ No |  |
|  | Large observation window with magnetic door seal |  | ☐ Yes ☐ No |  |
|  | Built-in lighting device, |  | ☐ Yes ☐ No |  |
|  | accurate measurement of the amount of oxygen that is consumed by microorganisms. |  | ☐ Yes ☐ No |  |
|  | can be used to optimize wastewater treatment processes, which is essential for protecting the environment and public health. |  | ☐ Yes ☐ No |  |
|  | Cooling system with high quality compressor, Refrigerant: Fluorine-free |  | ☐ Yes ☐ No |  |
|  | Over temperature alarm function |  | ☐ Yes ☐ No |  |
|  | Air Circulation Fan: Uniform temperature distribution depends on an air circulation fan, which distributes air across the chamber. |  | ☐ Yes ☐ No |  |
|  | Removable Shelves or Trays: Shelves and trays provide adjustable placement areas that suit different sample sizes along with multiple sample quantities. |  | ☐ Yes ☐ No |  |
|  | Safety Thermostat: The safety thermostat functions as a protective system through its capability to stop heating element power when the temperature reaches predetermined threshold levels. |  | ☐ Yes ☐ No |  |
|  | Digital PID Temperature Controller: A digital PID Temperature Controller operates with precise functionality to control the heating and cooling systems at desired temperatures. |  | ☐ Yes ☐ No |  |
|  | Control Panel: Through the Control Panel, users can set all operational settings, including temperature exposure duration as well as other parameters. |  | ☐ Yes ☐ No |  |
|  | Door Handle with Lock and Key |  | ☐ Yes ☐ No |  |
|  | Caster wheels |  | ☐ Yes ☐ No |  |
|  | The system maintains precise temperature control by allowing heat and cool cycles to work together. |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  | ☐ Yes ☐ No |  |
|  | UPS to Maintain 20-30 Min full Operation of the device |  | ☐ Yes ☐ No |  |
| 9 | Stirrer with Hot Plate |  |  |  |
|  | Name of Manufacturer | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | size of the medium-hot plates: 250 * 250 mm |  | ☐ Yes ☐ No |  |
|  | TEMPERATURE |  | ☐ Yes ☐ No |  |
|  | 30 ⁰ to 370 ⁰C |  | ☐ Yes ☐ No |  |
|  | "Thermostat Setting Increments 1°C steps |  | ☐ Yes ☐ No |  |
|  | Thermostat Working Accuracy ± 1°C" |  | ☐ Yes ☐ No |  |
|  | Digital temperature control to allow precise temperature control |  | ☐ Yes ☐ No |  |
|  | Body Type: Stainless Steel |  | ☐ Yes ☐ No |  |
|  | timer:0 to 99 min |  | ☐ Yes ☐ No |  |
|  | Safety controller for overheating protection |  | ☐ Yes ☐ No |  |
|  | Temperature uniformity |  | ☐ Yes ☐ No |  |
|  | Digital Display of Speed |  | ☐ Yes ☐ No |  |
|  | User Programmable Auto Shuts off |  | ☐ Yes ☐ No |  |
|  | Speed Adjustment: Rpm; 250-1500. |  | ☐ Yes ☐ No |  |
|  | Comes with at least 30 magnetic bars and Transparent protective cover included. |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |
| 10 | Incubator (CO2) |  |  |  |
|  | Name of Manufacturer | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | General Description |  | ☐ Yes ☐ No |  |
|  | Type: Bench-top laboratory incubator. |  | ☐ Yes ☐ No |  |
|  | Capacity: Approximately 150 liters internal chamber volume. |  | ☐ Yes ☐ No |  |
|  | Applications: Suitable for microbiology, clinical diagnostics, pharmaceutical testing and general lab incubation. |  | ☐ Yes ☐ No |  |
|  | HEAT CONDUCTION: Water-Jacketed |  | ☐ Yes ☐ No |  |
|  | Construction and Design |  | ☐ Yes ☐ No |  |
|  | Exterior: Powder-coated steel or corrosion-resistant metal housing. |  | ☐ Yes ☐ No |  |
|  | Interior: Mirror-finish stainless steel (SS 304 or better) with rounded internal corners for easy cleaning. |  | ☐ Yes ☐ No |  |
|  | UNIT COMPLETE WITH 2 CO2 CYLINDERS AND REGULATORS |  | ☐ Yes ☐ No |  |
|  | Sterlization Cycle |  | ☐ Yes ☐ No |  |
|  | Access Port |  | ☐ Yes ☐ No |  |
|  | Shelving: |  | ☐ Yes ☐ No |  |
|  | Minimum of 4 stainless steel shelves or drawers, adjustable or removable, capable of supporting various sample types. |  | ☐ Yes ☐ No |  |
|  | Door: |  | ☐ Yes ☐ No |  |
|  | Outer insulated metal door with magnetic gasket for sealing. |  | ☐ Yes ☐ No |  |
|  | Internal tempered glass window or secondary inner door for viewing without opening the main door. |  | ☐ Yes ☐ No |  |
|  | Heating System: |  | ☐ Yes ☐ No |  |
|  | Resistive heating elements combined with fan-assisted air circulation (forced convection) for uniform temperature and faster recovery after door opening. |  | ☐ Yes ☐ No |  |
|  | Temperature Range: Ambient +5°C to +70°C. |  | ☐ Yes ☐ No |  |
|  | Temperature Accuracy: ±0.1°C. |  | ☐ Yes ☐ No |  |
|  | Uniformity: ±0.25°C throughout the chamber. |  | ☐ Yes ☐ No |  |
|  | Sensor: High-precision PT-100 or equivalent. |  | ☐ Yes ☐ No |  |
|  | Controller: Microprocessor-based PID controller. |  | ☐ Yes ☐ No |  |
|  | HUMIDITY |  | ☐ Yes ☐ No |  |
|  | % Relative Humidity:Ambient to 95 @ 37 ⁰C |  | ☐ Yes ☐ No |  |
|  | Alarm :Hi/Lo |  | ☐ Yes ☐ No |  |
|  | Water Reservoir :Outside Chamber |  | ☐ Yes ☐ No |  |
|  | GAS |  | ☐ Yes ☐ No |  |
|  | Type (Co2) |  | ☐ Yes ☐ No |  |
|  | Solid State IR CO2 Sensor |  | ☐ Yes ☐ No |  |
|  | CO2 Control from 0-20% |  | ☐ Yes ☐ No |  |
|  | Programmability and Control Interface |  | ☐ Yes ☐ No |  |
|  | Display: Large LCD or touch screen showing set and actual parameters. |  | ☐ Yes ☐ No |  |
|  | Programmable Modes: |  | ☐ Yes ☐ No |  |
|  | Capability to save and execute multiple temperature-time programs |  | ☐ Yes ☐ No |  |
|  | Control Functions: |  | ☐ Yes ☐ No |  |
|  | real-time data display, and locked modes to prevent tampering. |  | ☐ Yes ☐ No |  |
|  | Data Logging: |  | ☐ Yes ☐ No |  |
|  | Internal memory for storing logs |  | ☐ Yes ☐ No |  |
|  | Safety and Alarm Features |  | ☐ Yes ☐ No |  |
|  | Over/under temperature, door open, sensor error, CO2, RH |  | ☐ Yes ☐ No |  |
|  | Safety Mechanisms |  | ☐ Yes ☐ No |  |
|  | Independent safety thermostat. |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
| 11 | Biological safety Cabinet class II, Type B2 |  |  |  |
|  | Name of Manufacturer | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Internal working area dimensions approx. (W x D x H, in mm): 1500 x 600 x 750, permissible deviation in width +/-10%, in depth and height +/-15% |  | ☐ Yes ☐ No |  |
|  | to exhaust all inflow and downflow air through ducting system to atmosphere |  | ☐ Yes ☐ No |  |
|  | Stainless steel cabinet housing with antibacterial coating |  | ☐ Yes ☐ No |  |
|  | Front window can be opened for easy cleaning |  | ☐ Yes ☐ No |  |
|  | In-built timer: available |  | ☐ Yes ☐ No |  |
|  | Display with air flow volume and speed values, HEPA filter operating time in %, total operating time, UV and fluorescent lamps operating time values |  | ☐ Yes ☐ No |  |
|  | Audible and visual alarm system for compromised airflow |  | ☐ Yes ☐ No |  |
|  | Automatic UV lamp turn-off when the front window is opened |  | ☐ Yes ☐ No |  |
|  | Audio and visual notifications for air flow speed, system availability, front window position, alarms for lamp replacement or maintenance |  | ☐ Yes ☐ No |  |
|  | Compatible stand for BSC installation, featuring casters and a position-locking function |  | ☐ Yes ☐ No |  |
|  | Average inflow velocity of approx. 0,5 m/s (100 fpm) or better |  | ☐ Yes ☐ No |  |
|  | No air re-circulation |  | ☐ Yes ☐ No |  |
|  | Air flow adjustment damper |  | ☐ Yes ☐ No |  |
|  | At least 304 type stainless steel working surface |  | ☐ Yes ☐ No |  |
|  | Removable work tray (work surface) for easy cleaning and decontamination |  | ☐ Yes ☐ No |  |
|  | at least 2 waterproof EU standard sockets |  | ☐ Yes ☐ No |  |
|  | Noise level maximum 58 dB (A) in standard mode |  | ☐ Yes ☐ No |  |
|  | Lighting in working area at least 1250 lux |  | ☐ Yes ☐ No |  |
|  | Front glass thickness at least 6 mm |  | ☐ Yes ☐ No |  |
|  | Negative pressure design |  | ☐ Yes ☐ No |  |
|  | Variable speed built-in blower of approximately 1/3 HP |  | ☐ Yes ☐ No |  |
|  | The Cabinet to be supplied with the following: |  | ☐ Yes ☐ No |  |
|  | HP totally enclosed fan-cooled remote exhaust blower. |  | ☐ Yes ☐ No |  |
|  | Built-in standard service fixtures |  | ☐ Yes ☐ No |  |
|  | Duct connection kit |  | ☐ Yes ☐ No |  |
|  | Two HEPA (99.99% efficient) filters |  | ☐ Yes ☐ No |  |
|  | Built-in Glare-free fluorescent lighting |  | ☐ Yes ☐ No |  |
|  | Built-in UV lamp |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | "A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | Compliant with IEC 60601 or equivalent." |  | ☐ Yes ☐ No |  |
| 12 | Fume Hood |  |  |  |
|  | Name of Manufacturer | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  |  |
|  | Designed and manufactured according to EN14175 or ASHRAE 110, with valid certificates confirming compliance with these standards. |  | ☐ Yes ☐ No |  |
|  | Ducted fume hood to exhaust hazardous or toxic chemical fumes, vapors, gasses and dust to protect lab environment and operator |  | ☐ Yes ☐ No |  |
|  | Floor mounted |  | ☐ Yes ☐ No |  |
|  | A ductwork type exhaust |  | ☐ Yes ☐ No |  |
|  | Built-in exhaust blower |  | ☐ Yes ☐ No |  |
|  | CONSTRUCTION |  | ☐ Yes ☐ No |  |
|  | Interior: Stainless Steel |  | ☐ Yes ☐ No |  |
|  | Exterior: Steel frame Epoxy coated |  | ☐ Yes ☐ No |  |
|  | Air flow design: Average inflow velocity adjustable at least from 0,3 to 0,6 m/s, or better |  | ☐ Yes ☐ No |  |
|  | Outflow air outlet into a duct through Carbone filter properly installed near the exhaust port of fume hood |  | ☐ Yes ☐ No |  |
|  | Pre-set baffle to provide a uniform face velocity pattern across the sash opening of the hood |  | ☐ Yes ☐ No |  |
|  | Vertical-rising tempered safety glass sash with 70cm sash opening when fully open |  | ☐ Yes ☐ No |  |
|  | Adjustable motorized transparent front panel with secure gasket and pedal control of its position |  | ☐ Yes ☐ No |  |
|  | Control Panel with LCD or LED screen for time, set-up and monitoring air velocity and temperature via sensors |  | ☐ Yes ☐ No |  |
|  | Light sources providing at least 800 lux light intensity |  | ☐ Yes ☐ No |  |
|  | Audible and visual alarm system for at least compromised airflow |  | ☐ Yes ☐ No |  |
|  | Internal working area dimensions approx. (W x D x H, in mm): 1500 x 700 x 900 |  | ☐ Yes ☐ No |  |
|  | Duplex Electrical Outlet to be provided built-in inside the chamber. |  | ☐ Yes ☐ No |  |
|  | Air flow monitor, Carbon filters for applications involving xylene and organic solvents. |  | ☐ Yes ☐ No |  |
|  | Total exhaust: must cover minimum 1000 CFM (100 FPM) at Fully Opened Sash |  | ☐ Yes ☐ No |  |
|  | Maximum Available Space by design (with base cabinet): (WxDxH) 1200x800x2400 mm |  | ☐ Yes ☐ No |  |
|  | To be supplied complete with duct connection kit, Ducting work, and connection from fume hood to outside/roofline to be carried out by the Electromechanical Contractor |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |
| 13 | Laminar Air Flow |  |  |  |
|  | Name of Manufacturer | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Designed and manufactured according to EN14175 or ASHRAE 110, with valid certificates confirming compliance with these standards. |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Floor-mounted type on a base stand with free leg space for the operator. |  | ☐ Yes ☐ No |  |
|  | Maintain minimum airflow velocity of 100 ft/ min (0.5 m/s) through the work access opening |  | ☐ Yes ☐ No |  |
|  | Interior working area aligned with the HEPA filters to provide minimized air turbulence. |  | ☐ Yes ☐ No |  |
|  | Filter efficiency of at least 99,99% MPPS at 0.3 μm particle size |  | ☐ Yes ☐ No |  |
|  | Provided with removable stainless-steel work trays. |  | ☐ Yes ☐ No |  |
|  | Front door is made of acrylic material sliding sash having counterbalance. |  | ☐ Yes ☐ No |  |
|  | Height-adjustable angled transparent front panel made of UV Resistant glass or better material with secure gasket and seal when in place |  | ☐ Yes ☐ No |  |
|  | Mobile and Height Adjustable castor base stand to be provided with cabinet. |  | ☐ Yes ☐ No |  |
|  | Control panel for time, set-up and monitoring air velocity via sensors |  | ☐ Yes ☐ No |  |
|  | Balanced Centrifugal Blower type |  | ☐ Yes ☐ No |  |
|  | Fluorescent glare-free lighting positioned at front side and UV Light with timer and interlock facility |  | ☐ Yes ☐ No |  |
|  | Include UV light timer from 0 to 60 minutes, for precise decontamination cycle control. |  | ☐ Yes ☐ No |  |
|  | Audible and visual alarm system for at least compromised airflow |  | ☐ Yes ☐ No |  |
|  | Air cleanliness level in the working area shall comply with at least class 5 as per ISO 14644-1 or an equivalent standard |  | ☐ Yes ☐ No |  |
|  | Duplex Electrical Outlet to be provided built-in inside the chamber. |  | ☐ Yes ☐ No |  |
|  | Interior is constructed from 304 grade stainless steel with no gaps or joints. (Or better) |  | ☐ Yes ☐ No |  |
|  | A built-in, low vibration and variable speed blower to direct filtered air across the working area. |  | ☐ Yes ☐ No |  |
|  | Maximum Available Space by design (with base cabinet): (WxDxH) 1400x800x2200mm |  | ☐ Yes ☐ No |  |
|  | To be supplied with 2 spare filters. |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |
| 14 | Bunsen Burner |  |  |  |
|  | Name of Manufacturer | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  |  |
|  | made of non-corrosive materials/steel and chromated brass. |  | ☐ Yes ☐ No |  |
|  | with nickel-plated air regulation |  | ☐ Yes ☐ No |  |
|  | with needle valve for selection of gas type |  | ☐ Yes ☐ No |  |
|  | suitable for propane and natural gas |  | ☐ Yes ☐ No |  |
|  | "Propane gas, gas pressure range 47.5-57.5 mbar |  | ☐ Yes ☐ No |  |
|  | (pressure regulator to be included)" |  | ☐ Yes ☐ No |  |
|  | Natural gas, gas pressure range 18-25 mbar |  | ☐ Yes ☐ No |  |
|  | tempreture Propane gas 1245 °C/Natural gas 1250 °C |  | ☐ Yes ☐ No |  |
|  | practical control of flame from very slow heating to rapid boiling with an adjustable air and gas regulation ring. Easy and precise flame adjustment is guaranteed by a screw driven gas control knob. |  | ☐ Yes ☐ No |  |
| 15 | Hoods, Canopy |  |  |  |
|  | Name of Manufacturer | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Designed to vent large lab system that are odor generating devise. |  | ☐ Yes ☐ No |  |
|  | Protecting both the user and the environment. |  | ☐ Yes ☐ No |  |
|  | The unit shall be wall mounted and or Ceiling mounted |  | ☐ Yes ☐ No |  |
|  | The canopy should be made of stainless-steel material |  | ☐ Yes ☐ No |  |
|  | The unit shall be equipped with fluorescent light, light switch, and blower switch. |  | ☐ Yes ☐ No |  |
|  | Provide complete set of accessories for mounting purpose. |  | ☐ Yes ☐ No |  |
|  | Remote blower should be included with unit. The supplier should coordinate with main contractor for the installation of remote blower |  | ☐ Yes ☐ No |  |
|  | Total air exhaust approximately: 700 CFM (with no re-circulation) |  | ☐ Yes ☐ No |  |
|  | Approximate Dimensions -free space needed (WxDxH): 1400X800X450 mm |  | ☐ Yes ☐ No |  |
|  | Shop-drawings must be coordinated with MEP services provided by contractor at site |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |

Delivery requirements –– Comparative Data Table

| UNOPS Requirements |  | Is quotation compliant? Bidder to complete | Details Bidder to complete |
| --- | --- | --- | --- |
| Delivery time | Bidder shall deliver the goods [120 calendar days] after Contract signature. | ☐ Yes ☐ No | Insert details |
| Delivery place and Incoterms rules | Delivery at Place DAP to Hospitals and Institutes mentioned in Bills of quantity tables- Form B, in case that the site is not ready, the supplier agrees to store the equipment in their warehouse until the site is prepared to receive the equipment. Incoterms rules as per Incoterms 2020. | ☐ Yes ☐ No | Insert details |
| Consignee details | This Shall include the Installation and Training requirements. Consignee will be the UNOPS and MoH representative | ☐ Yes ☐ No | Insert details |
| UNOPS Right to vary requirements | At the time the Contract is awarded, UNOPS reserves the right to vary the quantity of the goods and associated services specified above, provided this does not exceed +/- [20%] , without any change in the unit prices or other terms and conditions of the RFQ. | ☐ Yes ☐ No | Insert details |

00ax: +45 45 33 75 01

The offered goods and related services (if applicable) are in accordance with the required specifications and requirements specified in Section II: Schedule of Requirements.

☐ Yes   ☐ No

ANY DEVIATION MUST BE LISTED BELOW:

_________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

Name	: _____________________________________________________________

Title	: _____________________________________________________________

Date	: _____________________________________________________________

Signature	: _____________________________________________________________

If subcontracting is anticipated in the performance by the contractor, Procurement Official should insert the following:

List of subcontractors or suppliers

Bidder must identify the names of all subcontractors/suppliers who will be providing good/services under this contract and the type of work being subcontracted, if applicable.

- _[Full legal name and address of subcontractors]___________

- _________________________________________________

(C)_________________________________________________

I, the undersigned, certify that I am duly authorized by [insert full name of bidder] to sign this bid and bind [insert full name of bidder] should UNOPS accept this bid:

Name	: _____________________________________________________________

Title	: _____________________________________________________________

Date	: _____________________________________________________________

Signature	: _____________________________________________________________

# Form D: Previous Experience Form

RFQ reference no: [Insert UNOPS tender reference number]

Name of Bidder: [insert name of Bidder]

| Description of services/goods | Country | Total amount of Contract | Contract Identification and Title and Contact details of Client (Name, Address, telephone, email, fax) | Year project was undertaken |
| --- | --- | --- | --- | --- |

Name	: _____________________________________________________________

Title	: _____________________________________________________________

Date	: _____________________________________________________________

Signature	: _____________________________________________________________

# Form E: Self Disclosure Form

RFQ reference no: [insert RFQ reference No.]

Name of Bidder: [insert name of Bidder]

| Ref. | Critical area | Response | Comments |
| --- | --- | --- | --- |
| 1 | Has the entity or individual declared bankruptcy, or been involved in bankruptcy or receivership proceedings, or is there any judgment or pending legal action against them, which could impair operations in the foreseeable future; | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 2 | Has the entity or individual been found (or is awaiting a judgment or administrative decision) in breach of their obligations relating to the payment of taxes or social security contributions? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 3 | Has the entity or individual been found (or is awaiting a judgment or administrative decision) guilty of misconduct for the violation of applicable laws, regulations or ethical standards of the profession to which they belong? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 4 | Has the entity or individual engaged, or attempted to engage, in any Proscribed Practices in the past? For the purposes of this provision, “Proscribed Practices” are defined in the UNOPS Operational Instruction on Vendor Sanctions and include the following: - A corrupt practice: the offering, giving, receiving, or soliciting, directly or indirectly, anything of value to influence improperly the actions of another party; - A fraudulent practice: any act or omission, including a misrepresentation, that knowingly or recklessly misleads, or attempts to mislead, a party to obtain a financial or other benefit or to avoid an obligation; - A coercive practice: an act or omission that impairs or harms, or threatens to impair or harm, directly or indirectly, any party or the property of the party to improperly influence the actions of a party; - A collusive practice: an arrangement between two or more parties designed to achieve an improper purpose, including influencing improperly the actions of another party; - An unethical practice: Conduct or behaviour that is contrary to the conflict of interest, gifts and hospitality, post-employment provisions or other published requirements of doing business with UNOPS; and - Obstruction: Acts or omissions by a Vendor that prevent or hinder UNOPS from investigating instances of possible Proscribed Practices. | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 5 | Has the entity or individual unduly obtained, or attempted to unduly obtain, any confidential information in connection with this procurement process and any agreement that may be awarded as a result of this process? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 6 | Does the entity or individual have a conflict of interest, as provided in the Instructions to bidders, Article 3 “Bidder eligibility”, that may prevent them from entering into an agreement with UNOPS, or that may in any way jeopardize their performance in carrying out their obligations under the agreement? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 7 | Has the entity or individual been found (or is awaiting a judgment or administrative decision) to be involved in any of the following: - Fraudulent practice; - Corrupt practice; - Affiliation with a criminal organization; - Money laundering; - Terrorist financing; - Child labour; or - Human trafficking? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 8 | Has the entity or individual had significant performance issues in the past that led to the early termination of a legal commitment or the application of damages or other contractual penalties, or that were discovered following inspections, audits or investigations? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 9 | Has the entity or individual committed any financial irregularity, resulting from an act or omission or the inclusion of an unjustified item of expenditure outside a contractual engagement that had the effect of prejudicing their use of public funds? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 10 | Has the entity or individual been found (or is awaiting a judgment or administrative decision) to have created an entity in a different jurisdiction with the intent of circumventing fiscal, social or any other legal obligations in the jurisdiction of their registered office, central administration or principal place of business? | Select | [If the response is ‘Yes’, provide an explanation here.] |

I, the undersigned, hereby declare that the information provided is true and correct. I also understand that any willful dishonesty may result in the refusal of this submission. I understand that engagement in any of the above critical areas may, at the discretion of UNOPS, automatically result in the exclusion from this procurement process.

| Name: |  |
| --- | --- |
| Title |  |
| Date: | Sep 10, 2024 |
| Signature |  |

[Stamp this form with the official stamp of the bidder]

Form F: Joint Venture Partner Information Form

[The Bidder shall fill in this Form in accordance with the instructions indicated below].

RFQ reference no: [insert RFQ reference No.]

Name of Bidder: [insert name of bidder]

Date: [insert submission date]

To be completed and returned with your Bid if the Bid is submitted as a Joint Venture/Consortium/Association.

| JV / Consortium/ Association Information |  |
| --- | --- |
| Name | [complete] |
| Names of each partner and contact information (address, telephone numbers, fax numbers, e-mail address) | [complete] |
| Name of leading partner (with authority to bind the JV, Consortium, Association during the Bidding process and, in the event a Contract is awarded, during contract execution) | [complete] |
| Proposed proportion of responsibilities between partners (in %) with indication of the type of the goods/services to be delivered by each | [complete] |

Signatures of all partners of the JV:

We hereby confirm that if the contract is awarded, all parties of the Joint Venture/Consortium/Association shall be jointly and severally liable to UNOPS for the fulfillment of the provisions of the Contract.

Name of partner: ________________________		Name of partner: _________________________

Signature: ______________________________		Signature: _______________________________

Date: _______________________				Date: ________________________

Name of partner: ________________________		Name of partner: _________________________

Signature: ______________________________		Signature: _______________________________

Date: _______________________				Date: ________________________

Form G: Manufacturer’s Authorization Form

A letter issued by the manufacturer authorizing the applicant to participate in this particular RFQ must be submitted with the bid in the format provided in this Form.

To be eligible for delivery of goods, the bidder must be either the manufacturer of the offered goods or a sole representative of the manufacturer to the United Nations. Should offers for a particular make and model be received from more than one appointed representative, UNOPS reserves the right to select only one.

RFQ reference no: [Insert UNOPS tender reference number]

Name of Bidder: [insert name of bidder]

Date: [insert submission date]

To: [bidder to insert]

WHEREAS

We [insert complete name of manufacturer], who are official manufacturers of [insert type of goods manufactured], having factories at [insert full address of manufacturer’s factories], do hereby authorize [insert complete name of bidder] to submit a bid the purpose of which is to provide the following goods, manufactured by us [insert name and or brief description of the goods], and to subsequently negotiate and sign the contract.

We hereby extend our full guarantee and warranty in accordance with Clause 13 of the General Conditions for Goods, with respect to the goods offered by the above firm.

Signed: [insert signature(s) of authorized representative(s) of the manufacturer]

Name: [insert complete name(s) of authorized representative(s) of the manufacturer]

Title: [insert title]

Dated on ____________ day of __________________, _______ [insert date of signing]