Section III: Returnable Bidding Forms

eSourcing reference:  [RFQ/2026/62289]

Note to Bidders: The following returnable forms are part of this RFQ and must be completed and returned by bidders as part of their quotation. Instructions to complete each Form are highlighted in blue in each Form. Please complete the Returnable Bidding Forms as instructed and return them as part of your quotation by uploading them against their specific Document Checklist in the UNOPS eSourcing system.

This Section comprises the following Returnable Bidding Forms:

- Form A: Quotation Submission Form

- Form B: Price Schedule Form

- Form C: Technical Quotation Form

- Form D: Previous Experience Form

- Form E: Self Disclosure Form

- Form F: Joint Venture Partner Information Form

- Form G: Manufacturer’s authorization form

Form A: Quotation submission form

Bidders are requested to complete this form, sign it and return it as part of their bid submission. The bidder shall fill in this form in accordance with the instructions indicated. No alterations to its format shall be permitted and no substitutions shall be accepted.

Date: [Insert submission date]

Subject: Quotation for the supply of [Insert a brief description of goods/services] in [Name of country/city], RFQ Case No. [Insert RFQ ref number], dated [insert date]

We, the undersigned, declare that:

  - We offer to supply in conformity with the bidding documents, including the UNOPS General Conditions of Contract;

  - Our quotation shall be valid for the period of time of [insert number of days which shall not be less than the specified in the Tender Particulars section, Period of Validity of Quotations] from the date fixed for the submission deadline as set out in the RFQ, and it shall remain binding upon us and may be accepted at any time before the expiration of that period;

  - We have no conflict of interest in any activity that would put it, if selected for this assignment, in a conflict of interest with UNOPS [If you have any actual or potential conflict of interest as defined in Article 3 of Section I: Instructions to Bidders, please disclose it here];;

  - Our firm confirms that the offeror and sub-contractors have not been associated, or had been involved in any way, directly or indirectly, with the preparation of the design, terms of references and/or other documents used as a part of this solicitation;

  - Our firm, its affiliates or subsidiaries—including any subcontractors or suppliers for any part of the Contract—has not been declared ineligible by UNOPS, nor is included in the suspended/ineligibility list of the UN/PD, other UN Agencies, the UN Security Council, and the World Bank, in accordance with Instructions to Bidders Article 3, Eligibility;

  - We embrace the UN Supplier Code of Conduct and adhere to the principles of the UN Global Compact;

  - We have not declared bankruptcy, are not involved in bankruptcy or receivership proceedings, and there is no judgment or pending legal action against them that could impair their operations in the foreseeable future;

  - We have not offered and will not offer fees, gifts and/or favours of kind in exchange for this RFQ and will not engage in any such activity during the performance of any Contract awarded.

I, the undersigned, certify that I am duly authorized by [insert full name of bidder] to sign this quotation and bind [insert full name of bidder] should UNOPS accept this quotation:

Name: [complete]

Title: [complete]

Date: [complete]

Signature: _____________________________________________________________

Provide the name and contact information for the primary contact from your company for this quotation:

Name: [complete]

Title: [complete]

Email address: [complete]

Telephone: [complete]

# Form B: Price Schedule Form

Bidders shall fill in this Price Schedule Form in accordance with the instructions indicated.

**Please fill in all your prices in the attached excel sheet and send it as (excel book) along with the below table**

Pricing Methodology:

- In addition to DAP price bidder to add a breakdown of prices: cost of goods cleared for export (FCA) separate from transport and insurance.

- The Supplier shall, within fourteen (14) days of Contract signature, provide a contingency plan for alternative routing (e.g., land corridors or bypass ports) should the primary maritime route be restricted.

- To confirm whether the primary components and finished goods are sourced from regions not currently subject to maritime blockades or significant transit disruptions in the Strait of Hormuz or Red Sea (Geographic Origin Declaration).

RFQ reference no: [insert RFQ reference No.]

| Currency | USD |
| --- | --- |

Bills of quantity Lot# 3: Supply, installation and commissioning of Core lab & B.B+PATH lab medical equipment

| Item No | Description | Qty | Locations and Distributions | Unit Price DAP | Total Price DAP |
| --- | --- | --- | --- | --- | --- |
|  | Automated Coagulation Device | 2 | National Central Lab, Lahj | insert | insert |
|  | Full AUTOMATIC ELECTROPHORESIS | 2 | National Central Lab, Lahj | insert | insert |
|  | Fully Automated Chemistry Analyzer | 1 | National Central Lab, Lahj:1 | insert | insert |
|  | Full Automated Chemiluminescence Immunology Analyzer (Eliza) | 2 | National Central Lab, Lahj Radfan Hospital - Lahj:1 | insert | insert |
|  | Blood Gas analyzer | 3 | Ataq TFC, Shabwah: 1 Al-Mansoura TFC, Aden:1 National Central Lab, Lahj: 1 | insert | insert |
|  | Automatic Hematology Analyzer (CBC) | 5 | National Central Lab, Lahj: 1 Ataq TFC, Shabwah: 1 Al-Mansoura TFC, Aden:1 Hais TFC, Al Hudayda: 1 Radfan Hospital - Lahj:1 | insert | insert |
|  | Semi Auto Chemistry "Spectrophotometer" | 5 | National Central Lab, Lahj:1 Hais TFC, Al Hudayda: 1 Ataq TFC, Shabwah: 1 Al-Mansoura TFC, Aden:1 Radfan Hospital - Lahj:1 | insert | insert |
|  | Laboratory Incubator ~100 L | 1 | National Central Lab, Lahj | insert | insert |
|  | Laboratory Water Bath | 1 | National Central Lab, Lahj | insert | insert |
|  | Hemoglobin Meter | 6 | National Central Lab, Lahj | insert | insert |
|  | Electrolyte Machine | 2 | National Central Lab, Lahj: 1 Radfan Hospital - Lahj: 1 | insert | insert |
|  | Blood Tubes Mixer | 3 | National Central Lab, Lahj:2 Hais TFC, Al Hudayda: 1 | insert | insert |
|  | Laboratory Hotplate | 2 | Ataq TFC, Shabwah:1 Al-Mansoura TFC, Aden:1 | insert | insert |
|  | Slide warmer | 2 | Ataq TFC, Shabwah: 1 Al-Mansoura TFC, Aden:1 | insert | insert |
|  | Hematology, Slide Stainers | 1 | Dr. Ameen Nasher, Aden Building A:1 | insert | insert |
|  | Microbiology Slide Stainers | 1 | Dr. Ameen Nasher, Aden Building A:1 | insert | insert |
|  | Stethoscope, Binaural (Adult & children) | 1 | National Central Lab | insert | insert |
|  | Micropipettes (Fix) set (Different sizes) | 4 | Ataq TFC: 2 Al-Mansoura TFC, :2 | insert | insert |
|  | Micropipettes (Variable) (Different sizes) | 4 | Ataq TFC: 2 Al-Mansoura TFC, :2 | insert | insert |
|  | Platelets Incubator with Agitator | 3 | National Central Lab, Lahj: 1 Ataq TFC, Shabwah: 1 Al-Mansoura TFC, Aden:1 | insert | insert |
|  | Plasma Bath, Thawing | 1 | Radfan Hospital - Lahj:1 | insert | insert |
|  | Blood Collection Monitor | 2 | National Central Lab, Lahj: 1 Radfan Hospital - Lahj:1 | insert | insert |
|  | Blood Bag Tube Sealer | 1 | Radfan Hospital - Lahj:1 | insert | insert |
|  | Plasma Extrator | 1 | Radfan Hospital - Lahj:1 | insert | insert |
|  | Freezer, Laboratory, 1-door | 1 | Radfan Hospital - Lahj:1 | insert | insert |
|  | Plasma Refrigerator | 1 | National Central Lab, Lahj | insert | insert |
|  | Blood Bank Refrigerator | 1 | National Central Lab, Lahj | insert | insert |
|  | LAB Refrigerator | 3 | National Central Lab, Lahj: 1 Hais TFC, Al Hudayda: 1 Radfan Hospital - Lahj:1 | insert | insert |
|  | Blood Bag Refrigerated Centrifuge | 2 | National Central Lab, Lahj: 1 Radfan Hospital - Lahj:1 | insert | insert |
|  | Platelet Incubators, Benchtop | 1 | Radfan Hospital - Lahj:1 | insert | insert |
|  | Blood Warmer | 4 | Ataq TFC, Shabwah:2 Al-Mansoura TFC, Aden:2 | insert | insert |
|  | Warmer, Fluid/ Blood, Portable | 2 | Al Hezma, Mariab: 1 Radfan Hospital - Lahj:1 | insert | insert |
|  | Warmer, Infant, Care System | 1 | Al Hezma, Mariab | insert | insert |
|  | Blender, Gas, Air/Oxygen | 1 | Al Hezma, Mariab | insert | insert |
|  | Medical Refrigerator | 2 | Ataq TFC, Shabwah:1 Al-Mansoura TFC, Aden:1 | insert | insert |
| Total Price |  |  |  |  | insert |

Payment terms 30 days accepted: ☐ Yes

Bidder’s discount for accelerated payment: ____% of total firm price for each calendar day less than thirty (30) days

List of subcontractors or suppliers

Bidder must identify the names of all subcontractors/suppliers who will be providing good/services under this Contract and the type of work being subcontracted, if applicable.

- _[Full legal name and address of subcontractors]___________

- _________________________________________________

- _________________________________________________

I, the undersigned, certify that I am duly authorized by [insert full name of Bidder] to sign this quotation and bind [insert full name of Bidder] should UNOPS accept this quotation:

Name	: _____________________________________________________________

Title	: _____________________________________________________________

Date	: _____________________________________________________________

Signature	: _____________________________________________________________

Form C: Technical Quotation Form

RFQ reference no: [insert RFQ reference No.]

Name of Bidder: [insert name of Bidder]

Bidders are required to complete the Comparative Data Tables included in Section II: Schedule of Requirements to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.

Technical specifications for goods – Comparative Data Table

| Item No | UNOPS minimum technical requirements | Qty |  | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete |
| --- | --- | --- | --- | --- | --- |
| 1 | Automated Coagulation Device |  |  |  |  |
|  | Name of Manufacturer |  | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Bench-top coagulation analyzer. |  |  | ☐ Yes ☐ No |  |
|  | Usage: Suitable for medium-throughput clinical laboratories. |  |  | ☐ Yes ☐ No |  |
|  | Test Menu |  |  | ☐ Yes ☐ No |  |
|  | PT (Prothrombin Time) |  |  | ☐ Yes ☐ No |  |
|  | APTT (Activated Partial Thromboplastin Time) |  |  | ☐ Yes ☐ No |  |
|  | TT (Thrombin Time) |  |  | ☐ Yes ☐ No |  |
|  | Fibrinogen |  |  | ☐ Yes ☐ No |  |
|  | D-Dimer |  |  | ☐ Yes ☐ No |  |
|  | Antithrombin III |  |  | ☐ Yes ☐ No |  |
|  | FDP (Fibrin Degradation Products) |  |  | ☐ Yes ☐ No |  |
|  | Throughput: Minimum 60 PT tests/hour. |  |  | ☐ Yes ☐ No |  |
|  | System Configuration |  |  | ☐ Yes ☐ No |  |
|  | Channels: At least 4–8 independent detection channels. |  |  | ☐ Yes ☐ No |  |
|  | Reagent Positions: ≥10 reagent positions with onboard cooling (2–8°C). |  |  | ☐ Yes ☐ No |  |
|  | Sample Positions: ≥20 sample positions with STAT sample priority access. |  |  | ☐ Yes ☐ No |  |
|  | Cuvettes: Automatic loading of disposable cuvettes; minimum 500 onboard capacity. |  |  | ☐ Yes ☐ No |  |
|  | Pipetting: Automatic pipettor with clot detection and liquid-level sensing. |  |  | ☐ Yes ☐ No |  |
|  | Incubation & Reaction |  |  | ☐ Yes ☐ No |  |
|  | Reaction Temperature: Maintained at 37°C ± 0.5°C. |  |  | ☐ Yes ☐ No |  |
|  | Incubation Stations: At least 8 incubating positions for reagent/sample mixing. |  |  | ☐ Yes ☐ No |  |
|  | User Interface & Software |  |  | ☐ Yes ☐ No |  |
|  | Control: Integrated touchscreen display or external PC. |  |  | ☐ Yes ☐ No |  |
|  | Software: User-friendly software with automatic result calculation and flagging of abnormal results. |  |  | ☐ Yes ☐ No |  |
|  | Memory: Minimum 50,000 test results stored. |  |  | ☐ Yes ☐ No |  |
|  | Connectivity: USB port, LAN, RS232 for LIS/HIS integration. |  |  | ☐ Yes ☐ No |  |
|  | Printer and Output |  |  | ☐ Yes ☐ No |  |
|  | Printer: Integrated or external thermal printer for result printouts. |  |  | ☐ Yes ☐ No |  |
|  | Barcode Reader: Built-in or external barcode scanner for sample and reagent ID. |  |  | ☐ Yes ☐ No |  |
|  | Reagents & Calibration |  |  | ☐ Yes ☐ No |  |
|  | Open vial stability ≥ 7 days |  |  | ☐ Yes ☐ No |  |
|  | Refrigerated onboard storage |  |  | ☐ Yes ☐ No |  |
|  | Calibration: Automatic or semi-automatic, with QC functions. |  |  | ☐ Yes ☐ No |  |
|  | QC Levels: At least 2 levels of control per analyte; programmable schedules. |  |  | ☐ Yes ☐ No |  |
|  | REAGENT HANDLING |  |  | ☐ Yes ☐ No |  |
|  | Reagent Type |  |  | ☐ Yes ☐ No |  |
|  | Number of Reagents onboard / Tests per Container |  |  | ☐ Yes ☐ No |  |
|  | Cost /Test ($) |  |  | ☐ Yes ☐ No |  |
|  | Reagent Validity |  |  | ☐ Yes ☐ No |  |
|  | Safety and Maintenance |  |  | ☐ Yes ☐ No |  |
|  | Automatic waste disposal and liquid waste level sensor. |  |  | ☐ Yes ☐ No |  |
|  | Daily, weekly, and monthly maintenance routines should be simple and operator-guided. |  |  | ☐ Yes ☐ No |  |
|  | Self-diagnostics: Automatic error detection and notification system. |  |  | ☐ Yes ☐ No |  |
|  | Accessories and Consumables |  |  | ☐ Yes ☐ No |  |
|  | Start-up kit with reagents and cuvettes (minimum 500 tests) |  |  | ☐ Yes ☐ No |  |
|  | External printer (if not integrated) |  |  | ☐ Yes ☐ No |  |
|  | Barcode scanner (if not integrated) |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  |  | ☐ Yes ☐ No |  |
| 2 | Full AUTOMATIC ELECTROPHORESIS |  |  |  |  |
|  | Name of Manufacturer |  | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Type: Fully automated gel-based electrophoresis analyzer. |  |  | ☐ Yes ☐ No |  |
|  | Applications: Serum protein electrophoresis (SPE), urine protein, HbA1c, immunofixation, and similar clinical protocols. |  |  | ☐ Yes ☐ No |  |
|  | Automation: Complete automation from sample application through electrophoresis, staining, destaining, drying, and scanning. |  |  | ☐ Yes ☐ No |  |
|  | Processing Chamber: |  |  | ☐ Yes ☐ No |  |
|  | Enclosed and shielded chamber with transparent viewing window. |  |  | ☐ Yes ☐ No |  |
|  | Safety interlock: Chamber door must lock during active runs. |  |  | ☐ Yes ☐ No |  |
|  | Robotic Arm: |  |  | ☐ Yes ☐ No |  |
|  | Motorized arm for automatic gel loading, buffer dispensing, reagent handling, and stage transitions. |  |  | ☐ Yes ☐ No |  |
|  | Fully automated workflow from sample loading to result output. |  |  | ☐ Yes ☐ No |  |
|  | High-Voltage Power Supply |  |  | ☐ Yes ☐ No |  |
|  | Fully built-in, shielded DC power supply. |  |  | ☐ Yes ☐ No |  |
|  | Voltage Range: Adjustable from 0–1000V (or higher if required). |  |  | ☐ Yes ☐ No |  |
|  | Current Output: Up to 100 mA. |  |  | ☐ Yes ☐ No |  |
|  | Voltage Stability: ≤±1% during operation. |  |  | ☐ Yes ☐ No |  |
|  | Safety Features: |  |  | ☐ Yes ☐ No |  |
|  | Automatic shutdown for overvoltage or fault. |  |  | ☐ Yes ☐ No |  |
|  | Leakage protection and circuit insulation. |  |  | ☐ Yes ☐ No |  |
|  | Operator lockout for unauthorized voltage modification. |  |  | ☐ Yes ☐ No |  |
|  | Sample and Reagent Handling |  |  | ☐ Yes ☐ No |  |
|  | Sample Capacity: Minimum 20 samples per batch. |  |  | ☐ Yes ☐ No |  |
|  | Sample Types: Serum, plasma, urine, CSF, and whole blood (HbA1c). |  |  | ☐ Yes ☐ No |  |
|  | Identification: Integrated or external barcode scanner for samples and reagents. |  |  | ☐ Yes ☐ No |  |
|  | Reagent Management: |  |  | ☐ Yes ☐ No |  |
|  | Onboard reagent storage with automated dispensing. |  |  | ☐ Yes ☐ No |  |
|  | Reagent tracking (lot number, expiry, volume status). |  |  | ☐ Yes ☐ No |  |
|  | Electrophoresis and Gel Handling |  |  | ☐ Yes ☐ No |  |
|  | Gel Type: Compatible with precast agarose or cellulose acetate gels. |  |  | ☐ Yes ☐ No |  |
|  | Automation of Steps: |  |  | ☐ Yes ☐ No |  |
|  | Electrophoresis migration |  |  | ☐ Yes ☐ No |  |
|  | Staining and destaining |  |  | ☐ Yes ☐ No |  |
|  | Drying |  |  | ☐ Yes ☐ No |  |
|  | Gel transport to scanner |  |  | ☐ Yes ☐ No |  |
|  | Temperature Regulation: Programmable migration temperature (e.g., 20–30°C ±1°C). |  |  | ☐ Yes ☐ No |  |
|  | Detection and Analysis |  |  | ☐ Yes ☐ No |  |
|  | Detection System: Integrated high-resolution scanner (≥600 dpi). |  |  | ☐ Yes ☐ No |  |
|  | Software: |  |  | ☐ Yes ☐ No |  |
|  | Automatic band/zoning detection and quantification. |  |  | ☐ Yes ☐ No |  |
|  | Graphical and numerical output (percentage and absolute). |  |  | ☐ Yes ☐ No |  |
|  | Reference library for SPE, HbA1c, and IFE profiles. |  |  | ☐ Yes ☐ No |  |
|  | Manual override and review capability. |  |  | ☐ Yes ☐ No |  |
|  | Export in PDF and LIS-compatible formats (HL7, ASTM). |  |  | ☐ Yes ☐ No |  |
|  | Control Interface and Data Management |  |  | ☐ Yes ☐ No |  |
|  | User Interface: Touchscreen control panel or external PC. |  |  | ☐ Yes ☐ No |  |
|  | Storage Capacity: ≥50,000 test records. |  |  | ☐ Yes ☐ No |  |
|  | Connectivity: USB, LAN, RS232; full LIS/HIS integration. |  |  | ☐ Yes ☐ No |  |
|  | Reporting and Output |  |  | ☐ Yes ☐ No |  |
|  | Printing: Integrated or external printer (thermal or inkjet). |  |  | ☐ Yes ☐ No |  |
|  | Reports: Customizable with electropherogram, quantitative results, and interpretation space. |  |  | ☐ Yes ☐ No |  |
|  | Accessories and Consumables |  |  | ☐ Yes ☐ No |  |
|  | Initial start-up kit (≥500 tests) |  |  | ☐ Yes ☐ No |  |
|  | Barcode scanner (if not built-in) |  |  | ☐ Yes ☐ No |  |
|  | External printer (if not integrated) |  |  | ☐ Yes ☐ No |  |
|  | PC and licensed software (if applicable) |  |  | ☐ Yes ☐ No |  |
|  | Compatible UPS (≥30 minutes backup) |  |  | ☐ Yes ☐ No |  |
|  | Gels, reagents, stains, and consumables must be available from the supplier for at least 7 years. |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  |  | ☐ Yes ☐ No |  |
| 3 | Fully Automated CHEMISTRY/ IMMUNOASSAY ANALYZER |  |  |  |  |
|  | Name of Manufacturer |  | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Fully automated immunochemistry (chemistry and immunoassay) analyzer in one system |  |  | ☐ Yes ☐ No |  |
|  | Including all necessary accessories, reagents, calibrators, controls, and software required for operation, installation, training, and maintenance. |  |  | ☐ Yes ☐ No |  |
|  | Random access with STAT capability |  |  | ☐ Yes ☐ No |  |
|  | Operating Type: Batch / Random Access / Continous |  |  | ☐ Yes ☐ No |  |
|  | Closed system |  |  | ☐ Yes ☐ No |  |
|  | Must have a throughput capability around Throughput: Minimum 150 tests/hour (without ISE) (Chemistry & electrolytes) and around 25 (or better) tests/hour for Immunoassay. |  |  | ☐ Yes ☐ No |  |
|  | Supplier should provide with list of all assays that can be performed on the system. |  |  | ☐ Yes ☐ No |  |
|  | Hemolysis / Turbidity detection |  |  | ☐ Yes ☐ No |  |
|  | Reproducibility: CV ≤2% for most routine parameters |  |  | ☐ Yes ☐ No |  |
|  | Linearity: As per IFCC/CLSI standards |  |  | ☐ Yes ☐ No |  |
|  | Carry-over: ≤0.1% |  |  | ☐ Yes ☐ No |  |
|  | REAGENT HANDLING |  |  | ☐ Yes ☐ No |  |
|  | Reagent Tray Temperature:2 to 8⁰C |  |  | ☐ Yes ☐ No |  |
|  | "Reagent volume: 100–500 µL or lower Barcode reading for reagents" |  |  | ☐ Yes ☐ No |  |
|  | Specify the number of reagents in the refrigerated positions |  |  | ☐ Yes ☐ No |  |
|  | Specify Reagent type (ready to use, calibration frequency, stability …) |  |  | ☐ Yes ☐ No |  |
|  | Must provide on-board management of reagents (random placement, on-board stability and calibration tracking) |  |  | ☐ Yes ☐ No |  |
|  | SAMPLE HANDLING |  |  | ☐ Yes ☐ No |  |
|  | Must accept the following sample Types: Serum, plasma, urine, CSF |  |  | ☐ Yes ☐ No |  |
|  | Clot detection or aspiration error alert system |  |  | ☐ Yes ☐ No |  |
|  | Calibration and Quality Control |  |  | ☐ Yes ☐ No |  |
|  | Multi-point calibration |  |  | ☐ Yes ☐ No |  |
|  | On-board QC management |  |  | ☐ Yes ☐ No |  |
|  | QC graphing and Westgard rules support |  |  | ☐ Yes ☐ No |  |
|  | Auto re-run and auto dilution functions |  |  | ☐ Yes ☐ No |  |
|  | Tube-to-tube carryover must be less than 0.1 ppm or better |  |  | ☐ Yes ☐ No |  |
|  | Capability for STAT samples interrupt |  |  | ☐ Yes ☐ No |  |
|  | Specify number of STAT samples that can be loaded. |  |  | ☐ Yes ☐ No |  |
|  | Specify the number of Maximum number of sample capacity. |  |  | ☐ Yes ☐ No |  |
|  | Specify the aspiration system and the aspiration volume |  |  | ☐ Yes ☐ No |  |
|  | Sample Tubes:Different Size Tubes / Sample Cups / Microcontainer |  |  | ☐ Yes ☐ No |  |
|  | Specify the type of the light source and all filters installed (wavelength..) |  |  | ☐ Yes ☐ No |  |
|  | Sample Bar Codes |  |  | ☐ Yes ☐ No |  |
|  | User Interface and Software |  |  | ☐ Yes ☐ No |  |
|  | LCD MONITOR |  |  | ☐ Yes ☐ No |  |
|  | Touchscreen or computer-based control |  |  | ☐ Yes ☐ No |  |
|  | High resolution printer for printing of patient report, QC and calibration curves, etc. |  |  | ☐ Yes ☐ No |  |
|  | User configurable printing format |  |  | ☐ Yes ☐ No |  |
|  | LIS/HIS connectivity (HL7 compatible) |  |  | ☐ Yes ☐ No |  |
|  | Data storage: Minimum 50,000 results |  |  | ☐ Yes ☐ No |  |
|  | control software that control and monitor all the process with Online Error Code Help |  |  | ☐ Yes ☐ No |  |
|  | CALIBRATION / QC |  |  | ☐ Yes ☐ No |  |
|  | Store up around 5 calibration curves per test |  |  | ☐ Yes ☐ No |  |
|  | Reports to include patient’s demographic information |  |  | ☐ Yes ☐ No |  |
|  | water purification system specifically designed for use with the offered unit (water quality and capacity) shall be included with the offer, Consumables for such system shall be priced |  |  | ☐ Yes ☐ No |  |
|  | Accessories & Consumables |  |  | ☐ Yes ☐ No |  |
|  | External Printer |  |  | ☐ Yes ☐ No |  |
|  | UPS (Uninterruptible Power Supply) with minimum 30 minutes backup |  |  | ☐ Yes ☐ No |  |
|  | Based on average 50 samples/day |  |  | ☐ Yes ☐ No |  |
|  | Accessories, Reagents & consumables for the Tests: To be priced separately |  |  | ☐ Yes ☐ No |  |
|  | Number of tests for the above priced reagents and consumables. |  |  | ☐ Yes ☐ No |  |
|  | Cost /Test |  |  | ☐ Yes ☐ No |  |
|  | Cost/Test will be taken into consideration during the calculation of the total cost of ownership. |  |  | ☐ Yes ☐ No |  |
|  | Starter pack of reagents, calibrators, and controls for 3 months |  |  | ☐ Yes ☐ No |  |
|  | User training for lab personnel |  |  | ☐ Yes ☐ No |  |
|  | Local service support must be available (response time <24 hours) |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | "A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions |  |  | ☐ Yes ☐ No |  |
| 4 | Full Automated Immunoassay Analyzer |  |  |  |  |
|  | Name of Manufacturer |  | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Principle: Should be based on the principle of chemiluminescence / Electrochemiluminescence Immunoassay technology with very high sensitivity and linearity. |  |  | ☐ Yes ☐ No |  |
|  | Test Types: The test types will include hormones, tumor markers, antibodies to infectious diseases, auto immune diseases, allergy, cardiac, nutrition markers, markers of bone metabolism, reproductive markers and Triple markers. |  |  | ☐ Yes ☐ No |  |
|  | Photometer Light source: Halogen Reading mode: OD (8 channel head) Dynamic range (measurement range): –0.100 to 3 000 OD Accuracy: 0 – 2 000 OD 2.5 % Linearity: 0 – 2 000 OD 1 % Detection: Photodiode Filters: Up to 8 positions (delivered with 450 and 620 nm) |  |  | ☐ Yes ☐ No |  |
|  | Pipetting/Processing System Processing time samples: Approx. 16 minutes Processing time reagents: Approx. 4 minutes Min/max volumes: 10 μl to 300 μl (with 300 μl tip) / or 301 to 1000 μl (with 1100 μl tip) Precision: < 5 % CV at 25 μl; < 2.5 % at 100 μl Features: Tip detection, mixing, multi-dispensing |  |  | ☐ Yes ☐ No |  |
|  | Incubation Capacity: 4 independently controlled chambers; 4 independent ambient stackers Mixing capability: Mixing frequency adjustable Temperature range: Ambient + 7 to 45 °C Accuracy: ± 2°C Uniformity: – 1.5 °C; + 1 °C |  |  | ☐ Yes ☐ No |  |
|  | Safety features > Clot and bubble detection > Password protection > Automatic reagent level checks > Barcode reader for primary tubes > Capacity level sensing > Advanced level-error logging and execution |  |  | ☐ Yes ☐ No |  |
|  | Continuous loading > New tests/plates can be loaded at any time |  |  | ☐ Yes ☐ No |  |
|  | Provision for running emergency (stat) tests. |  |  |  |  |
|  | Test Throughput: Fully automated with a throughput of 60 tests or more per hour. |  |  | ☐ Yes ☐ No |  |
|  | Barcode identification of samples and reagents |  |  | ☐ Yes ☐ No |  |
|  | System should have high resolution touch screen. |  |  | ☐ Yes ☐ No |  |
|  | Compatible with Laboratory information system for easy reporting. |  |  | ☐ Yes ☐ No |  |
|  | All Supplied Reagents Should Have a Minimum Shelf life of 12 Months. |  |  | ☐ Yes ☐ No |  |
|  | Based on average 50 samples/day |  |  | ☐ Yes ☐ No |  |
|  | Accessories, Reagents & consumables for the Tests: To be priced separately |  |  | ☐ Yes ☐ No |  |
|  | Number of tests for the above priced reagents and consumables. |  |  | ☐ Yes ☐ No |  |
|  | Cost /Test |  |  | ☐ Yes ☐ No |  |
|  | Cost/Test will be taken into consideration during the calculation of the total cost of ownership. |  |  | ☐ Yes ☐ No |  |
|  | Accessories: |  |  | ☐ Yes ☐ No |  |
|  | Pc Computer With Software and Color printer etc |  |  | ☐ Yes ☐ No |  |
|  | All consumables required for installation and standardization of system to be given free of cost |  |  | ☐ Yes ☐ No |  |
|  | Suitable Online UPS with maintenance free batteries for minimum one-hour back-up should be supplied with the system |  |  | ☐ Yes ☐ No |  |
|  | water purification system specifically designed for use with the offered unit (water quality and capacity) shall be included with the offer, Consumables for such system shall be priced |  |  | ☐ Yes ☐ No |  |
|  | Any software or database updates should be done free of cost by the firm, during the life of the equipment, as and when it is released by the manufacturer. |  |  | ☐ Yes ☐ No |  |
|  | Free One Kit for all Tests. |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
| 5 | Blood Gas analyzer |  |  |  |  |
|  | Name of Manufacturer |  | 3 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | TESTS AVAILABLE |  |  | ☐ Yes ☐ No |  |
|  | Blood Gases Module: pCO2, pO2, ph. |  |  | ☐ Yes ☐ No |  |
|  | Co-Oximetry Module: sO2, O2Hb, HHb, COHb, Met Hb. |  |  | ☐ Yes ☐ No |  |
|  | Derived |  |  | ☐ Yes ☐ No |  |
|  | a) Bicarbonate (HCO₃⁻) b) Standard HCO₃⁻ c) Base Excess (BE) d) Base Excess of Extra cellular fluid (BE-Ecf) e) Oxygen Saturation (SO2%) f) Anion Gap (AG) |  |  | ☐ Yes ☐ No |  |
|  | Accept heparinized whole blood, serum or plasma, arterial, mixed venous and capillary sample for analysis |  |  | ☐ Yes ☐ No |  |
|  | Sample Mode: Syringe, Capillary |  |  | ☐ Yes ☐ No |  |
|  | Indicative Sample size: less than 150 ul for Partial and Full Parameter Analysis. |  |  | ☐ Yes ☐ No |  |
|  | Time of Result: Time to result must never exceed 120 seconds. |  |  | ☐ Yes ☐ No |  |
|  | Throughput: Minimum 40 samples per hour or more. |  |  | ☐ Yes ☐ No |  |
|  | Liquid calibration method for calibration of all measured parameters without use of any gas tanks or humidifiers |  |  | ☐ Yes ☐ No |  |
|  | Clot Detection with Auto-Cleaning Feature is Demanded. |  |  | ☐ Yes ☐ No |  |
|  | Printer:. Include Built-In Printer. |  |  | ☐ Yes ☐ No |  |
|  | Barcode |  |  | ☐ Yes ☐ No |  |
|  | Feature barcode reading for reagents, sensors, QC, operator ID and sampler. |  |  | ☐ Yes ☐ No |  |
|  | Storge: Capacity to Store Minimum of 200 Patients Results or Better , 100 QC approximately . |  |  | ☐ Yes ☐ No |  |
|  | Audible and Visual Error Messages |  |  | ☐ Yes ☐ No |  |
|  | Screen display: Full Color LCD/TFT, (Touch Screen) |  |  | ☐ Yes ☐ No |  |
|  | Screen Display of Reagents Levels, Clotted Samples, and Bubbles |  |  | ☐ Yes ☐ No |  |
|  | Quilaty Control Mode: Include Automatic QC Mode and Maintenance Procedures. |  |  | ☐ Yes ☐ No |  |
|  | Interface: Capability for LIS Interfaces for Data Transfer and/or PC Compatibility for Data Export. |  |  | ☐ Yes ☐ No |  |
|  | Accessories: |  |  | ☐ Yes ☐ No |  |
|  | Shall be supplied Complete with The Necessary Start-Up Kit of Consumables and Free Reagents for Initial QC Testing, Calibration, System Validation and Training Activities. |  |  | ☐ Yes ☐ No |  |
|  | Bidder to Include Separate Quote for all Required Consumables and Reagents |  |  | ☐ Yes ☐ No |  |
|  | Number of tests for the above accessories and consumables |  |  | ☐ Yes ☐ No |  |
|  | Number of days for the above accessories and consumables |  |  | ☐ Yes ☐ No |  |
|  | CPT (Cost /Test )($) for all the above mentioned test. |  |  | ☐ Yes ☐ No |  |
|  | Machine Should be Equipped with UPS Online 2KVA. |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
|  | compliant with IEC 60601 or equivalent. |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  |  | ☐ Yes ☐ No |  |
| 6 | Automatic Hematology Analyzer (CBC) |  |  |  |  |
|  | Name of Manufacturer |  | 5 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Hematology analyzer for routine testing (whole blood samples) including: |  |  | ☐ Yes ☐ No |  |
|  | 28 PARAMETERS reported: WBC, RBC, HGB, HCT, MCV, MCH, MCHC,RDW-SD, RDW-CV , PLT, NEUT%, LYMPH %, MONO%, EOS %, BASO %, NEUT #, LYMPH #, MONO #,EOS #, BASO #, PDW, MPV, PCT, P-LCR, NRBC #, NRBC %, IG #, IG % Three Histograms – RBC, PLT and WBC |  |  | ☐ Yes ☐ No |  |
|  | RET% |  |  | ☐ Yes ☐ No |  |
|  | RBC & PLT-DC Detection |  |  | ☐ Yes ☐ No |  |
|  | Autoloader |  |  | ☐ Yes ☐ No |  |
|  | CYCLE TIME,30 -60 sec |  |  | ☐ Yes ☐ No |  |
|  | START-UP TIME(3 min) |  |  | ☐ Yes ☐ No |  |
|  | QC PROGRAMS:3 levels |  |  | ☐ Yes ☐ No |  |
|  | Built-in QC program for statistical purposes |  |  | ☐ Yes ☐ No |  |
|  | Specify fixed or adjustable (manual and automated) threshold for each parameter. |  |  | ☐ Yes ☐ No |  |
|  | CBC sample rate ~ 50 / hour (closed tube mode) |  |  | ☐ Yes ☐ No |  |
|  | CBCD sample rate ~ 30 / hour (closed tube mode) |  |  | ☐ Yes ☐ No |  |
|  | Sample volume should be around 200 μL in normal mode. Specify exact sample volume |  |  | ☐ Yes ☐ No |  |
|  | Data storage capacity (10,000 patients) |  |  | ☐ Yes ☐ No |  |
|  | High-resolution 8 "LCD screen for data entry, result viewing and all types of user interfacing |  |  | ☐ Yes ☐ No |  |
|  | High resolution printer for report printing |  |  | ☐ Yes ☐ No |  |
|  | Analysis modes: Manual and auto Sampler modes |  |  | ☐ Yes ☐ No |  |
|  | Reports to include histograms and scatter plots |  |  | ☐ Yes ☐ No |  |
|  | User configurable printing format |  |  | ☐ Yes ☐ No |  |
|  | Built-in automatic maintenance program, preferably zero user maintenance |  |  | ☐ Yes ☐ No |  |
|  | Audiovisual alarms for sample error, aspiration and clog errors, reagent detection error,internal malfunction, etc. shall be incorporated |  |  | ☐ Yes ☐ No |  |
|  | Internal troubleshooting software capability |  |  | ☐ Yes ☐ No |  |
|  | Computer connectivity for Bi-directional data, results and report transmission from and to the Hospital Information System |  |  | ☐ Yes ☐ No |  |
|  | Reagent and consumable items prices |  |  | ☐ Yes ☐ No |  |
|  | The supplier shall provide a table of calculated cost / test showing method of calculation and all included hidden costs such as wash solution, PM parts (filters, tubing, electrodes,etc.), sample cups, calibration kits, etc. |  |  | ☐ Yes ☐ No |  |
|  | At the time of Installation of equipment to consignee, company should supply reagents for 200 Test ( Diluent, Lyse, cleaner) |  |  | ☐ Yes ☐ No |  |
|  | Online UPS 1KVA |  |  | ☐ Yes ☐ No |  |
|  | Tube Component for Waste |  |  | ☐ Yes ☐ No |  |
|  | Tube Component for Diluent |  |  | ☐ Yes ☐ No |  |
|  | Tube Component for Lyse |  |  | ☐ Yes ☐ No |  |
|  | 20 L Bottle for Waste |  |  | ☐ Yes ☐ No |  |
|  | Set of Thermal Paper. |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
|  | compliant with IEC 60601 or equivalent. |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  |  | ☐ Yes ☐ No |  |
| 7 | Semi Auto Biochemistry (Spectrophotometer) |  |  |  |  |
|  | Name of Manufacturer |  | 5 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | spectrophotometer-based for use in clinical laboratories. |  |  | ☐ Yes ☐ No |  |
|  | Designed for qualitative and quantitative analysis of blood, serum, plasma, urine, and other biological fluids. |  |  | ☐ Yes ☐ No |  |
|  | Suitable for low- to medium-throughput laboratories (small hospitals, clinics, satellite labs). |  |  | ☐ Yes ☐ No |  |
|  | Open system (can use reagents from different manufacturers) |  |  | ☐ Yes ☐ No |  |
|  | Functional Requirements |  |  | ☐ Yes ☐ No |  |
|  | Measurement Principle: Photometric measurement based on absorbance. |  |  | ☐ Yes ☐ No |  |
|  | Use of filter-based or monochromator-based wavelength selection. |  |  | ☐ Yes ☐ No |  |
|  | Wavelength Range: At least 340 nm to 630 nm or wider. |  |  | ☐ Yes ☐ No |  |
|  | Wavelength Selection: Automatic or manual selection of at least 6 to 8 standard wavelengths (e.g., 340, 405, 505, 546, 578, 630 nm). |  |  | ☐ Yes ☐ No |  |
|  | Light Source: Long-life halogen or LED lamp. |  |  | ☐ Yes ☐ No |  |
|  | Throughput: Minimum of 100 tests per hour (manual sample and reagent pipetting). |  |  | ☐ Yes ☐ No |  |
|  | Cuvette Type: Compatible with reusable or disposable cuvettes. |  |  | ☐ Yes ☐ No |  |
|  | Minimum sample volume: 200–500 µL. |  |  | ☐ Yes ☐ No |  |
|  | Calibration |  |  | ☐ Yes ☐ No |  |
|  | Multi-point and single-point calibration modes. |  |  | ☐ Yes ☐ No |  |
|  | Automatic curve generation for standard calibration |  |  | ☐ Yes ☐ No |  |
|  | Quality Control |  |  | ☐ Yes ☐ No |  |
|  | Inbuilt QC program with Levey-Jennings plots and Westgard Rules support. |  |  | ☐ Yes ☐ No |  |
|  | User Interface |  |  | ☐ Yes ☐ No |  |
|  | Display: touchscreen display (minimum 7 inches) or color LCD. |  |  | ☐ Yes ☐ No |  |
|  | Memory capacity to store at least 100–200 programs/tests and 1000+ results. |  |  | ☐ Yes ☐ No |  |
|  | Sample and Reagent Management |  |  | ☐ Yes ☐ No |  |
|  | Manual pipetting of sample and reagent with clear guided instructions. |  |  | ☐ Yes ☐ No |  |
|  | Reaction time and absorbance readings programmable and customizable. |  |  | ☐ Yes ☐ No |  |
|  | Data Management |  |  | ☐ Yes ☐ No |  |
|  | Internal printer |  |  | ☐ Yes ☐ No |  |
|  | USB port for data backup and software updates. |  |  | ☐ Yes ☐ No |  |
|  | LIS/HIS connectivity (optional). |  |  | ☐ Yes ☐ No |  |
|  | Power Supply |  |  | ☐ Yes ☐ No |  |
|  | AC supply: 220–240 V, 50–60 Hz. |  |  | ☐ Yes ☐ No |  |
|  | Accessories and Consumables |  |  | ☐ Yes ☐ No |  |
|  | Starter pack of cuvettes (minimum 200 pcs). |  |  | ☐ Yes ☐ No |  |
|  | Starter pack of controls and calibration standards. |  |  | ☐ Yes ☐ No |  |
|  | Printer paper rolls |  |  | ☐ Yes ☐ No |  |
|  | barcode reader |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
|  | compliant with IEC 60601 or equivalent. |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  |  | ☐ Yes ☐ No |  |
| 8 | Laboratory Incubator 100 L |  |  |  |  |
|  | Name of Manufacturer |  | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Type: Bench-top or floor-standing laboratory incubator. |  |  | ☐ Yes ☐ No |  |
|  | Capacity: Approximately 100 liters internal chamber volume. |  |  | ☐ Yes ☐ No |  |
|  | Applications: Suitable for microbiology, clinical diagnostics, pharmaceutical testing, food safety, and general lab incubation. |  |  | ☐ Yes ☐ No |  |
|  | Construction and Design |  |  | ☐ Yes ☐ No |  |
|  | Exterior: Powder-coated steel or corrosion-resistant metal housing. |  |  | ☐ Yes ☐ No |  |
|  | Interior: Mirror-finish stainless steel (SS 304 or better) with rounded internal corners for easy cleaning. |  |  | ☐ Yes ☐ No |  |
|  | Class 100 HEPA filtered Air |  |  | ☐ Yes ☐ No |  |
|  | Sterilization Cycle |  |  |  |  |
|  | Access Port |  |  | ☐ Yes ☐ No |  |
|  | Shelving: |  |  | ☐ Yes ☐ No |  |
|  | Minimum of 3 stainless steel shelves or drawers, adjustable or removable, capable of supporting various sample types. |  |  | ☐ Yes ☐ No |  |
|  | Door: |  |  | ☐ Yes ☐ No |  |
|  | Outer insulated metal door with magnetic gasket for sealing. |  |  | ☐ Yes ☐ No |  |
|  | Internal tempered glass window or secondary inner door for viewing without opening the main door. |  |  | ☐ Yes ☐ No |  |
|  | Heating System: |  |  | ☐ Yes ☐ No |  |
|  | Resistive heating elements combined with fan-assisted air circulation (forced convection) for uniform temperature and faster recovery after door opening. |  |  | ☐ Yes ☐ No |  |
|  | Temperature Range: Ambient +5°C to +70°C. |  |  | ☐ Yes ☐ No |  |
|  | Temperature Accuracy: ±0.1°C. |  |  | ☐ Yes ☐ No |  |
|  | Uniformity: ±0.2°C throughout the chamber. |  |  | ☐ Yes ☐ No |  |
|  | Sensor: High-precision PT-100 or equivalent. |  |  | ☐ Yes ☐ No |  |
|  | Controller: Microprocessor-based PID controller. |  |  | ☐ Yes ☐ No |  |
|  | HUMIDITY |  |  | ☐ Yes ☐ No |  |
|  | % Relative Humidity: Ambient to 95 @ 37 ⁰C |  |  | ☐ Yes ☐ No |  |
|  | GAS |  |  | ☐ Yes ☐ No |  |
|  | Type (Co2) |  |  | ☐ Yes ☐ No |  |
|  | Solid State IR CO2 Sensor |  |  | ☐ Yes ☐ No |  |
|  | CO2 Control from 0-20% |  |  | ☐ Yes ☐ No |  |
|  | Programmability and Control Interface |  |  | ☐ Yes ☐ No |  |
|  | Display: Large LCD or touch screen showing set and actual parameters. |  |  | ☐ Yes ☐ No |  |
|  | Programmable Modes: |  |  | ☐ Yes ☐ No |  |
|  | Capability to save and execute multiple temperature-time programs |  |  | ☐ Yes ☐ No |  |
|  | Control Functions: |  |  | ☐ Yes ☐ No |  |
|  | real-time data display, and locked modes to prevent tampering. |  |  | ☐ Yes ☐ No |  |
|  | Data Logging: |  |  | ☐ Yes ☐ No |  |
|  | Internal memory for storing logs |  |  | ☐ Yes ☐ No |  |
|  | Safety and Alarm Features |  |  | ☐ Yes ☐ No |  |
|  | Over/under temperature, door open, sensor error, CO2, RH |  |  | ☐ Yes ☐ No |  |
|  | Safety Mechanisms |  |  | ☐ Yes ☐ No |  |
|  | Independent safety thermostat. |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
| 9 | Laboratory Water Bath |  |  |  |  |
|  | Name of Manufacturer |  | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Stainless steel tank well insulated. |  |  | ☐ Yes ☐ No |  |
|  | Heater and temperature sensors mounted underneath the tank and Low water level sensor |  |  | ☐ Yes ☐ No |  |
|  | Corrosion resistant and easy to clean with drain system. |  |  | ☐ Yes ☐ No |  |
|  | Complete with bath cover and handles |  |  | ☐ Yes ☐ No |  |
|  | Capacity: not less than 20 litres. |  |  | ☐ Yes ☐ No |  |
|  | Adjustable temperature from +25 Cº to + 99 Cº |  |  | ☐ Yes ☐ No |  |
|  | Water proof control panel. |  |  | ☐ Yes ☐ No |  |
|  | stainless steel Tube racks to hold 12x75mm and 16x100mm tubes, 30-50 tubes/rack |  |  | ☐ Yes ☐ No |  |
|  | Temperature to be adjusted in increments of 0.1 Cº |  |  | ☐ Yes ☐ No |  |
|  | Uniformity:+/- 0.2 C |  |  | ☐ Yes ☐ No |  |
|  | Digital temperature display. |  |  | ☐ Yes ☐ No |  |
|  | Safety cut-off thermostat |  |  | ☐ Yes ☐ No |  |
|  | Timer: About 60 Minutes Or Better. |  |  | ☐ Yes ☐ No |  |
|  | Accessories: |  |  | ☐ Yes ☐ No |  |
|  | Heater Spare Part |  |  | ☐ Yes ☐ No |  |
|  | Holder of Tubes: for 10 tubes |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
|  | compliant with IEC 60601 or equivalent. |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  |  | ☐ Yes ☐ No |  |
| 10 | Hemoglobin Meter |  |  |  |  |
|  | Name of Manufacturer |  | 6 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Type: Portable, point-of-care hemoglobin meter for quantitative measurement of hemoglobin concentration in whole blood. |  |  | ☐ Yes ☐ No |  |
|  | Application: For use in hospitals, clinics, blood donation centers, and field screening programs. |  |  | ☐ Yes ☐ No |  |
|  | Measurement Performance |  |  | ☐ Yes ☐ No |  |
|  | Measurement Range: 4.0 – 25.0 g/dL (or equivalent in mmol/L). |  |  | ☐ Yes ☐ No |  |
|  | Sample Type: Capillary, venous, or arterial whole blood. |  |  | ☐ Yes ☐ No |  |
|  | Sample Volume: ≤10 µL (microcuvette or test strip-based). |  |  | ☐ Yes ☐ No |  |
|  | Result Time: ≤60 seconds. |  |  | ☐ Yes ☐ No |  |
|  | Accuracy: ±0.3 g/dL or better compared to reference laboratory method. |  |  | ☐ Yes ☐ No |  |
|  | Repeatability: CV ≤2%. |  |  | ☐ Yes ☐ No |  |
|  | Display and Output |  |  | ☐ Yes ☐ No |  |
|  | LCD or digital display showing Hb result clearly in g/dL. |  |  | ☐ Yes ☐ No |  |
|  | Units: Option to display in g/dL and mmol/L (user selectable). |  |  | ☐ Yes ☐ No |  |
|  | Memory: Minimum 500 results stored internally with date/time stamp. |  |  | ☐ Yes ☐ No |  |
|  | Operation and Power |  |  | ☐ Yes ☐ No |  |
|  | Power Source: Rechargeable battery |  |  | ☐ Yes ☐ No |  |
|  | Battery Life: ≥1000 tests per charge or battery cycle. |  |  | ☐ Yes ☐ No |  |
|  | User Interface with 2–4 navigation buttons. |  |  | ☐ Yes ☐ No |  |
|  | Consumables and Accessories |  |  | ☐ Yes ☐ No |  |
|  | Test Method: Disposable microcuvettes or reagent strips (no need for external reagents). |  |  | ☐ Yes ☐ No |  |
|  | Starter Kit: Device should be supplied with, but not limited to; |  |  | ☐ Yes ☐ No |  |
|  | 1 device |  |  | ☐ Yes ☐ No |  |
|  | 1 box of 100 test cuvettes/strips |  |  | ☐ Yes ☐ No |  |
|  | 1 lancing device with 100 lancets |  |  | ☐ Yes ☐ No |  |
|  | Calibration/control solution |  |  | ☐ Yes ☐ No |  |
|  | Carrying case |  |  | ☐ Yes ☐ No |  |
|  | Power adapter or charger |  |  | ☐ Yes ☐ No |  |
|  | Quality and Calibration |  |  | ☐ Yes ☐ No |  |
|  | Auto Calibration |  |  | ☐ Yes ☐ No |  |
|  | Quality Control: Built-in self-checks or control solution support. |  |  | ☐ Yes ☐ No |  |
|  | Auto power-off after inactivity. |  |  | ☐ Yes ☐ No |  |
|  | Output and Printing |  |  | ☐ Yes ☐ No |  |
|  | The device should support external printing of results either: |  |  | ☐ Yes ☐ No |  |
|  | USB connection to an external portable thermal printer |  |  | ☐ Yes ☐ No |  |
|  | Printout should include: |  |  | ☐ Yes ☐ No |  |
|  | Hemoglobin result |  |  | ☐ Yes ☐ No |  |
|  | Date/time |  |  | ☐ Yes ☐ No |  |
|  | Patient/sample ID |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source If applicable |  |  | ☐ Yes ☐ No |  |
|  | compliant with IEC 60601 or equivalent. |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  |  | ☐ Yes ☐ No |  |
| 12 | Electrolyte Machine |  |  |  |  |
|  | Name of Manufacturer |  | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Method Used: ISE Measurement |  |  | ☐ Yes ☐ No |  |
|  | Open System for reagent and consumables |  |  | ☐ Yes ☐ No |  |
|  | Full Automatic operation |  |  | ☐ Yes ☐ No |  |
|  | Sample Type: Whole blood, plasma, serum, dialysate, diluted urine |  |  | ☐ Yes ☐ No |  |
|  | Sample Volume, µL 150 μlor less. |  |  | ☐ Yes ☐ No |  |
|  | Stat interrupt: Included |  |  | ☐ Yes ☐ No |  |
|  | Analysis Time, Sec: 90 Sec approximately. |  |  | ☐ Yes ☐ No |  |
|  | Throughput, Samples/hr: 40 Samples/Hrs Or Better. |  |  | ☐ Yes ☐ No |  |
|  | Test Menu: Na, K, Cl,Ca |  |  | ☐ Yes ☐ No |  |
|  | Assay Range : mmol/L |  |  | ☐ Yes ☐ No |  |
|  | Na+: 40-205 |  |  | ☐ Yes ☐ No |  |
|  | K+:1.5-15 |  |  | ☐ Yes ☐ No |  |
|  | Ca++:0.2-10 |  |  | ☐ Yes ☐ No |  |
|  | Calibration: |  |  | ☐ Yes ☐ No |  |
|  | Type: Automatic or On-demand |  |  | ☐ Yes ☐ No |  |
|  | Display: LED Or LCD Touchscreen |  |  | ☐ Yes ☐ No |  |
|  | Printer: Built-in thermal |  |  | ☐ Yes ☐ No |  |
|  | Bar-Code Reader |  |  | ☐ Yes ☐ No |  |
|  | Data Management: Quality control memory storage |  |  | ☐ Yes ☐ No |  |
|  | Interface: Rs-232C |  |  | ☐ Yes ☐ No |  |
|  | Equipped with : |  |  | ☐ Yes ☐ No |  |
|  | Solution Packs, Electrodes , conditioning Solutions, Cleaning Solutions, Calibrators & Controls |  |  | ☐ Yes ☐ No |  |
|  | Accessories: Accessories and Reagents to be included in the offer |  |  | ☐ Yes ☐ No |  |
|  | Cost per Test should be Provided Clearly and Electrodes (Na+, K+, and CI-) Price |  |  | ☐ Yes ☐ No |  |
|  | On-line UPS |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
|  | compliant with IEC 60601 or equivalent. |  |  | ☐ Yes ☐ No |  |
| 12 | Blood Tubes Mixer |  |  |  |  |
|  | Name of Manufacturer | 3 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Provided with Timer(1 Sec to 99 min) |  |  | ☐ Yes ☐ No |  |
|  | Number of Tube Rollers: 9-12 Rollers. |  |  | ☐ Yes ☐ No |  |
|  | Speed Variable: 1 to 70 rpm. |  |  | ☐ Yes ☐ No |  |
|  | Rolling Angle & Tilting Angle Rolling : ±3° . |  |  | ☐ Yes ☐ No |  |
|  | Digital readout and display. |  |  | ☐ Yes ☐ No |  |
|  | To be used in incubators up to 60C° and humidity up to 80%, or in cold rooms down to 4C°. |  |  | ☐ Yes ☐ No |  |
|  | Timed and continuous operation |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | Power Supply: AC 100~240V, 50/60Hz. |  |  | ☐ Yes ☐ No |  |
| 13 | Laboratory Hotplate |  |  |  |  |  |  |  |
|  | Name of Manufacturer | 2 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | size of the medium-hot plates: 250 * 250 mm |  |  | ☐ Yes ☐ No |  |
|  | Working temperatures up to 350 Celsius, |  |  | ☐ Yes ☐ No |  |
|  | plates made of aluminum |  |  | ☐ Yes ☐ No |  |
|  | Body Type: Stainless Steel |  |  | ☐ Yes ☐ No |  |
|  | Digital temperature control to allow precise temperature control |  |  | ☐ Yes ☐ No |  |
|  | timer:0 to 99 min |  |  | ☐ Yes ☐ No |  |
|  | Safety controller for overheating protection |  |  | ☐ Yes ☐ No |  |
|  | temperature uniformity |  |  | ☐ Yes ☐ No |  |
|  | Thermostat Setting Increments 1°C steps Thermostat Working Accuracy ± 1°C |  |  | ☐ Yes ☐ No |  |
|  | Digital Data Display |  |  | ☐ Yes ☐ No |  |
|  | User Programmable Auto Shuts off |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
|  | Compliant with IEC 60601 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  |  | ☐ Yes ☐ No |  |
| 14 | Slide Warmer |  |  |  |  |
|  | Name of Manufacturer | 2 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Equipment for drying paraffin tissue sections. |  |  | ☐ Yes ☐ No |  |
|  | Benchtop |  |  | ☐ Yes ☐ No |  |
|  | Thermal heater must ensure heat transfer over the entire surface |  |  | ☐ Yes ☐ No |  |
|  | Tempered glass lid which seals with integral gasket on heating base |  |  | ☐ Yes ☐ No |  |
|  | Anodized black surface to provide contrast with samples |  |  | ☐ Yes ☐ No |  |
|  | Should have digital temperature control and LED panel. |  |  | ☐ Yes ☐ No |  |
|  | Temperate setting by digital buttons and not by manual rotatory nob. |  |  | ☐ Yes ☐ No |  |
|  | Should hold upto 24 slides.. |  |  | ☐ Yes ☐ No |  |
|  | 5. Thermostatic temperature control from room temperature to 70°C and not more than 70°C. |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
|  | compliant with IEC 60601 or equivalent. |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  |  | ☐ Yes ☐ No |  |
| 15 | Slide Stainer ,Hematology |  |  |  |  |
|  | Name of Manufacturer | 1 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | APPLICATION: For staining microscope slides for Hematology and microbiology |  |  | ☐ Yes ☐ No |  |
|  | TYPE: Bench Top |  |  | ☐ Yes ☐ No |  |
|  | up to 12 slides in less than six minutes |  |  | ☐ Yes ☐ No |  |
|  | primary set up of four reagent vessels, one drying and one rinsing station |  |  | ☐ Yes ☐ No |  |
|  | Pumps Precise Volume of Reagent |  |  | ☐ Yes ☐ No |  |
|  | Accurate Mixing |  |  | ☐ Yes ☐ No |  |
|  | Uniform Stain & Buffer Delivery |  |  | ☐ Yes ☐ No |  |
|  | Alert for Low Level Reagent & Waste Full |  |  | ☐ Yes ☐ No |  |
|  | Stained Slide Dryer |  |  | ☐ Yes ☐ No |  |
|  | Cellular Application: Peripheral Blood & Bone Marrow Smears |  |  | ☐ Yes ☐ No |  |
|  | Slide Collection Drawer |  |  | ☐ Yes ☐ No |  |
|  | Waste Disposal |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  |  | ☐ Yes ☐ No |  |
| 16 | Slide Stainer , Microbiology |  |  |  |  |
|  | Name of Manufacturer | 1 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | APPLICATION: For staining microscope slides for Hematology and microbiology |  |  | ☐ Yes ☐ No |  |
|  | TYPE: Bench Top |  |  | ☐ Yes ☐ No |  |
|  | up to 12 slides in less than six minutes |  |  | ☐ Yes ☐ No |  |
|  | primary set up of four reagent vessels, one drying and one rinsing station |  |  | ☐ Yes ☐ No |  |
|  | Pumps Precise Volume of Reagent |  |  | ☐ Yes ☐ No |  |
|  | Accurate Mixing |  |  | ☐ Yes ☐ No |  |
|  | Uniform Stain & Buffer Delivery |  |  | ☐ Yes ☐ No |  |
|  | Alert for Low Level Reagent & Waste Full |  |  | ☐ Yes ☐ No |  |
|  | Stained Slide Dryer |  |  | ☐ Yes ☐ No |  |
|  | Cellular Application: Peripheral Blood & Bone Marrow Smears |  |  | ☐ Yes ☐ No |  |
|  | Slide Collection Drawer |  |  | ☐ Yes ☐ No |  |
|  | Waste Disposal |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  |  | ☐ Yes ☐ No |  |
| 17 | Stethoscope, Binaural (Adult & children) |  |  |  |  |
|  | Manufacturer | 1 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Double cup, dual-use (adult and pediatric auscultation) chest piece in stainless steel or chrome plated brass. |  |  | ☐ Yes ☐ No |  |
|  | Adult diaphragm 43-47mm; pediatric diaphragm 28-36mm. |  |  | ☐ Yes ☐ No |  |
|  | Y tube treated rubber or PVC |  |  | ☐ Yes ☐ No |  |
|  | Arms: Stainless Steel |  |  | ☐ Yes ☐ No |  |
|  | Removable Plastic Ear-Pieces. |  |  | ☐ Yes ☐ No |  |
|  | Latex-free |  |  | ☐ Yes ☐ No |  |
|  | Accessories: |  |  | ☐ Yes ☐ No |  |
|  | 4 X Ear caps made of elastomer rubber |  |  | ☐ Yes ☐ No |  |
|  | 1 X Storge Case |  |  | ☐ Yes ☐ No |  |
| 18 | Micropipettes (Fix) set (Different sizes) |  |  |  |  |
|  | Name of Manufacturer | 4 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Incorporated ejector mechanism |  |  | ☐ Yes ☐ No |  |
|  | Smooth running piston |  |  | ☐ Yes ☐ No |  |
|  | Sterilizable (without the need for recalibration) |  |  | ☐ Yes ☐ No |  |
|  | Simple calibration procedure (calibration tool to be included) |  |  | ☐ Yes ☐ No |  |
|  | Each unit shall have its individual calibration (conformity) certificate |  |  | ☐ Yes ☐ No |  |
|  | Ergonomic hand grip design with finger rest |  |  | ☐ Yes ☐ No |  |
|  | Rotary selector knob with clear volume display window |  |  | ☐ Yes ☐ No |  |
|  | Clear marking of volume range and color coding |  |  | ☐ Yes ☐ No |  |
|  | Safety label with plastic cover |  |  | ☐ Yes ☐ No |  |
|  | Blowout step to prevent capillary action in the tip for micro volumes (0.2 μl - 40 μl) |  |  | ☐ Yes ☐ No |  |
|  | The offer shall include a set of one single channel micropipettes for each of the volume specified below (manufacturer variations are accepted as long as the whole range is covered): |  |  | ☐ Yes ☐ No |  |
|  | Micropipette Fixed Volume Micro Pipettors in 12 volume sizes |  |  | ☐ Yes ☐ No |  |
|  | "Capacity: 5ul, 10ul, 20ul, 25ul, 50ul, 100ul, 200ul, 250ul, 500ul, 1000ul, 2000ul, 5000ul |  |  | ☐ Yes ☐ No |  |
|  | Tip compatibility: Universal |  |  | ☐ Yes ☐ No |  |
|  | Color code |  |  | ☐ Yes ☐ No |  |
|  | Volume display: ≥ 3 |  |  | ☐ Yes ☐ No |  |
|  | Volume adjustment lock |  |  | ☐ Yes ☐ No |  |
|  | Specify Accuracy of measurements |  |  | ☐ Yes ☐ No |  |
|  | Specify Precision (repeatability) |  |  | ☐ Yes ☐ No |  |
| 19 | Micropipettes (Variable) (Different sizes) |  |  |  |  |
|  | Name of Manufacturer | 4 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Incorporated ejector mechanism |  |  | ☐ Yes ☐ No |  |
|  | Smooth running piston |  |  | ☐ Yes ☐ No |  |
|  | Sterilizable (without the need for recalibration) |  |  | ☐ Yes ☐ No |  |
|  | Simple calibration procedure (calibration tool to be included) |  |  | ☐ Yes ☐ No |  |
|  | Each unit shall have its individual calibration (conformity) certificate |  |  | ☐ Yes ☐ No |  |
|  | Ergonomic hand grip design with finger rest |  |  | ☐ Yes ☐ No |  |
|  | Rotary selector knob with clear volume display window |  |  | ☐ Yes ☐ No |  |
|  | Clear marking of volume range and color coding |  |  | ☐ Yes ☐ No |  |
|  | Safety label with plastic cover |  |  | ☐ Yes ☐ No |  |
|  | Blowout step to prevent capillary action in the tip for micro volumes (0.2 μl - 40 μl) |  |  | ☐ Yes ☐ No |  |
|  | The offer shall include a set of one single channel micropipettes for each of the volume specified below (manufacturer variations are accepted as long as the whole range is covered): |  |  | ☐ Yes ☐ No |  |
|  | 0.2 μl - 2 μl |  |  | ☐ Yes ☐ No |  |
|  | 1 ml - 2.5 ml |  |  | ☐ Yes ☐ No |  |
|  | 0.5 μl - 10 μl |  |  | ☐ Yes ☐ No |  |
|  | 5 μl - 50 μl |  |  | ☐ Yes ☐ No |  |
|  | 10 μl- 100 μl |  |  | ☐ Yes ☐ No |  |
|  | 50 μl - 200 μl |  |  | ☐ Yes ☐ No |  |
|  | 100 μl - 1000 μl |  |  | ☐ Yes ☐ No |  |
|  | Multi-range accuracy parameters (within the specified volume range) for each unit shall be specified |  |  | ☐ Yes ☐ No |  |
|  | Maximum systematic error (inaccuracy) in absolute and relative volumes (μl and %). |  |  | ☐ Yes ☐ No |  |
|  | Maximum random error (imprecision) in SD and CV (SD in μl and CV in %) |  |  | ☐ Yes ☐ No |  |
|  | The offer shall include one shelf (circular) hanger, capable of holding the pipette set stated above |  |  | ☐ Yes ☐ No |  |
|  | its universal pipette tips, easy tip loading and ejecting. |  |  | ☐ Yes ☐ No |  |
|  | Color Coded |  |  | ☐ Yes ☐ No |  |
|  | Supplied with ≥250 tips with each unit |  |  | ☐ Yes ☐ No |  |
| 20 | Platelets Incubator with Agitator |  |  |  |  |
|  | Name of Manufacturer | 3 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | TEMPERATURE CONTROL WITH UNIFORM AIR CIRCULATION |  |  | ☐ Yes ☐ No |  |
|  | Designed to be used in conjunction with agitator |  |  | ☐ Yes ☐ No |  |
|  | Digital speed selection and timer |  |  | ☐ Yes ☐ No |  |
|  | Speed RPM:40-80 Approx. |  |  | ☐ Yes ☐ No |  |
|  | Time: 0 to 99 min |  |  | ☐ Yes ☐ No |  |
|  | Touch screen alarm and control system |  |  | ☐ Yes ☐ No |  |
|  | 7-day chart recorder |  |  | ☐ Yes ☐ No |  |
|  | Dual temperature display |  |  | ☐ Yes ☐ No |  |
|  | AUDIO VISUAL ALERT ALARMS |  |  | ☐ Yes ☐ No |  |
|  | Motion sensor |  |  | ☐ Yes ☐ No |  |
|  | Temperature alarm |  |  | ☐ Yes ☐ No |  |
|  | Door open alarm |  |  | ☐ Yes ☐ No |  |
|  | Remote alarm contacts |  |  | ☐ Yes ☐ No |  |
|  | Access port |  |  | ☐ Yes ☐ No |  |
|  | Capacity : 50 bags |  |  | ☐ Yes ☐ No |  |
|  | SS internal chamber and internal shelves |  |  | ☐ Yes ☐ No |  |
|  | Glass transparent door, with lock |  |  | ☐ Yes ☐ No |  |
|  | Online UPS |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  |  | ☐ Yes ☐ No |  |
| 21 | Plasma Bath, Thawing |  |  |  |  |
|  | Name of Manufacturer | 1 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Bench top with touch screen operation & LED display |  |  | ☐ Yes ☐ No |  |
|  | Temperature monitoring :Digital display |  |  | ☐ Yes ☐ No |  |
|  | Temperature accuracy: 0.1ºC |  |  | ☐ Yes ☐ No |  |
|  | Surface & Chamber Construction: Stainless Steel |  |  | ☐ Yes ☐ No |  |
|  | Thawing Capacity: 8 plasma bags |  |  | ☐ Yes ☐ No |  |
|  | Capability to thaws up to eight plasma units in around 25 minutes. |  |  | ☐ Yes ☐ No |  |
|  | Should have two separate basket assemblies with built-in fingers for securely holding the plasma bags of all sizes. |  |  | ☐ Yes ☐ No |  |
|  | Temperature and agitation control |  |  | ☐ Yes ☐ No |  |
|  | Should give an alarm when the plasma bags are thawed |  |  | ☐ Yes ☐ No |  |
|  | Provision for programmable time setting for length of thawing |  |  | ☐ Yes ☐ No |  |
|  | Should have digital timer clearly displaying the programmed set time or remaining cycle in minutes. |  |  | ☐ Yes ☐ No |  |
|  | Should have audio visual over-temperature alarm system |  |  | ☐ Yes ☐ No |  |
|  | Chamber volume:32 Liter |  |  | ☐ Yes ☐ No |  |
|  | Drain valve |  |  | ☐ Yes ☐ No |  |
|  | Temperature Adjustment Range °C: 20 to 40 |  |  | ☐ Yes ☐ No |  |
|  | Water Circulation |  |  | ☐ Yes ☐ No |  |
|  | Removable basket |  |  | ☐ Yes ☐ No |  |
|  | Thermal insulated lid |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  |  | ☐ Yes ☐ No |  |
| 22 | Blood Collection Monitor |  |  |  |  |
|  | Name of Manufacturer | 2 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | “micro controller” based Blood Collection Monitor system which collects, monitors and mixes the blood with anti-coagulant in the blood bag. It is light weight and easy to operate. |  |  | ☐ Yes ☐ No |  |
|  | Touch display with donor info, for clear visibility of status & prompt messages. |  |  | ☐ Yes ☐ No |  |
|  | Real time visual flow rate indicators. |  |  | ☐ Yes ☐ No |  |
|  | Audio and visual alert. |  |  | ☐ Yes ☐ No |  |
|  | Tarring of bag volume before collection. |  |  | ☐ Yes ☐ No |  |
|  | Automatically recall of last donation. |  |  | ☐ Yes ☐ No |  |
|  | • Continuous tray movement during collection or pause at 12-18 RPM. |  |  | ☐ Yes ☐ No |  |
|  | Automatically stops blood collection by clamping the tubing at the end of pre-selected set volume. |  |  | ☐ Yes ☐ No |  |
|  | Standard weight for calibration |  |  |  |  |
|  | Automatic release of clamp when bag is lifted at the end of pre-selected set volume. |  |  | ☐ Yes ☐ No |  |
|  | Suitable for all types of bags. |  |  | ☐ Yes ☐ No |  |
|  | Collection volume shows in ml/gms both. |  |  | ☐ Yes ☐ No |  |
|  | Automatic close of clamp during PAUSE and tray movement continuous. |  |  | ☐ Yes ☐ No |  |
|  | Facility of changing the pre-set volume in PAUSE mode. |  |  | ☐ Yes ☐ No |  |
|  | Single touch volume selection for standard collection of- 100 ml, 350 ml & 450 ml. |  |  | ☐ Yes ☐ No |  |
|  | Facility of emergency stop. |  |  | ☐ Yes ☐ No |  |
|  | Real Time Clock facility. |  |  | ☐ Yes ☐ No |  |
|  | Icon display facility for battery voltage level. LED indication for process identification. |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  |  | ☐ Yes ☐ No |  |
| 23 | Blood Bag Tube Sealer |  |  |  |  |
|  | Name of Manufacturer | 1 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | System should gently seal the tubing with no hemolysis. |  |  | ☐ Yes ☐ No |  |
|  | Capable of sealing tubes of various manufacturers of blood bag. |  |  | ☐ Yes ☐ No |  |
|  | Interchangeable Sealing Heads Available |  |  | ☐ Yes ☐ No |  |
|  | Sealing triggering should be automatic |  |  | ☐ Yes ☐ No |  |
|  | Built-in Safety Lock |  |  | ☐ Yes ☐ No |  |
|  | No warm-up time required. |  |  | ☐ Yes ☐ No |  |
|  | Detection of wet tube, leakage & sealing defect. |  |  | ☐ Yes ☐ No |  |
|  | Alarm in case of seal not complete |  |  | ☐ Yes ☐ No |  |
|  | Handheld configuration |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  |  | ☐ Yes ☐ No |  |
| 24 | Plasma Extractor |  |  |  |  |
|  | Name of Manufacturer | 1 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Compatible with all multiple blood bags. |  |  | ☐ Yes ☐ No |  |
|  | Type |  |  | ☐ Yes ☐ No |  |
|  | Squeeze plasma or other components from one bag to another |  |  | ☐ Yes ☐ No |  |
|  | Spring-loaded front panel to exert pressure on the bag to transfer the fluid until container is clamped off or emptied. |  |  | ☐ Yes ☐ No |  |
|  | Construction |  |  | ☐ Yes ☐ No |  |
|  | Transparent Plate for visual control red cells/plasma |  |  | ☐ Yes ☐ No |  |
| 25 | FREEZER, laboratory |  |  |  |  |
|  | Name of Manufacturer | 1 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | FREEZER, laboratory,. upright -20 deg.C |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | For specimen freezing to be used in the main lab |  |  | ☐ Yes ☐ No |  |
|  | Capacity shall be approximately 500 Liters Approx |  |  | ☐ Yes ☐ No |  |
|  | The temperature range must be between -15°C to -30°C or better |  |  | ☐ Yes ☐ No |  |
|  | A temperature monitor shall be included and integrated . The monitor will display chamber temperature to the tenth degree centigrade (°C). |  |  | ☐ Yes ☐ No |  |
|  | The temperature display shall be independent of the control circuitry so that malfunctioning control temperature sensing circuits can be detected. |  |  | ☐ Yes ☐ No |  |
|  | The display shall be visible in bright, dim, or sunlight, and the alarms shall be audible over background noise. |  |  | ☐ Yes ☐ No |  |
|  | Exterior Finish: Stainless Steel Front |  |  | ☐ Yes ☐ No |  |
|  | Interior Finish :Stainless Steel |  |  | ☐ Yes ☐ No |  |
|  | Heavy Duty Stainless Steel Pull out Drawers |  |  | ☐ Yes ☐ No |  |
|  | Temperature Recorder with Back up Battery |  |  | ☐ Yes ☐ No |  |
|  | Alarms |  |  | ☐ Yes ☐ No |  |
|  | High/low alarms that can be reprogrammed. High and low limits will be displayed (at least while they are being set). |  |  | ☐ Yes ☐ No |  |
|  | Audible and visual alarm to alert for power loss and/or temperature deviation beyond set limits. |  |  | ☐ Yes ☐ No |  |
|  | The laboratory freezer will have a method to test alarms. |  |  | ☐ Yes ☐ No |  |
|  | Audible alarm will not be defeatable, or will have a corresponding visual indicator during a temporary alarm silence period. |  |  | ☐ Yes ☐ No |  |
|  | The freezer shall have provision for contact to interface with the Building Management System (BMS) |  |  | ☐ Yes ☐ No |  |
|  | Construction: |  |  | ☐ Yes ☐ No |  |
|  | Door: Single, Solid Stainless Steel with Key Lock |  |  | ☐ Yes ☐ No |  |
|  | Locking casters, with Leveling |  |  | ☐ Yes ☐ No |  |
|  | At least five stainless steel racks. |  |  | ☐ Yes ☐ No |  |
|  | Sure-Seal self-closing door system |  |  | ☐ Yes ☐ No |  |
|  | Door Gasket: Magnetic |  |  | ☐ Yes ☐ No |  |
|  | Interior light. |  |  | ☐ Yes ☐ No |  |
|  | Enlarge Digital Display / Touch Screen |  |  | ☐ Yes ☐ No |  |
|  | Refrigeration System: |  |  | ☐ Yes ☐ No |  |
|  | Hermetically sealed, air-cooled compressor. |  |  | ☐ Yes ☐ No |  |
|  | Non-CFC refrigerant. |  |  | ☐ Yes ☐ No |  |
|  | Forced air circulation maintains chamber uniformity of +/-1°C and provides quick recovery. |  |  | ☐ Yes ☐ No |  |
|  | Interior fans shut down when door is opened. |  |  | ☐ Yes ☐ No |  |
|  | Automatic defrosting |  |  | ☐ Yes ☐ No |  |
|  | Automatic condensate evaporation system. |  |  | ☐ Yes ☐ No |  |
|  | Event log Recording with Date & Time |  |  | ☐ Yes ☐ No |  |
|  | Seven-day temperature chart recorder. |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  |  | ☐ Yes ☐ No |  |
| 26 | Plasma Refrigerator |  |  |  |  |
|  | Name of Manufacturer | 1 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Suitable for storing fresh plasma bags at 2–6°C in compliance with WHO and AABB standards. |  |  | ☐ Yes ☐ No |  |
|  | Storage Capacity |  |  | ☐ Yes ☐ No |  |
|  | Net internal volume: 350 liters |  |  | ☐ Yes ☐ No |  |
|  | Storage of at least 150 plasma bags (350 ml each) |  |  | ☐ Yes ☐ No |  |
|  | Modular or adjustable drawer or shelf system for plazma bag segregation |  |  | ☐ Yes ☐ No |  |
|  | plasma bag storage baskets or trays made of SS 304 or medical-grade plastic |  |  | ☐ Yes ☐ No |  |
|  | Temperature Control |  |  | ☐ Yes ☐ No |  |
|  | Temperature range: +2°C to +6°C |  |  | ☐ Yes ☐ No |  |
|  | Accuracy: ±1°C |  |  | ☐ Yes ☐ No |  |
|  | Uniform temperature distribution using forced air circulation |  |  | ☐ Yes ☐ No |  |
|  | Microprocessor-controlled |  |  | ☐ Yes ☐ No |  |
|  | Touchscreen control panel 7” with: Real-time temp display (current, min, max) Alarm history, trend graphs, and diagnostic alerts Password-protected access for settings |  |  | ☐ Yes ☐ No |  |
|  | Internal & External Construction |  |  | ☐ Yes ☐ No |  |
|  | External material: Powder-coated steel or stainless steel |  |  | ☐ Yes ☐ No |  |
|  | Internal chamber: High-grade stainless steel (SS 304 or better) |  |  | ☐ Yes ☐ No |  |
|  | Insulation: CFC-free, high-density polyurethane foam (≥75 mm) |  |  | ☐ Yes ☐ No |  |
|  | Door: Lockable, double-sealed, magnetic gasket |  |  | ☐ Yes ☐ No |  |
|  | Viewing window: double-glass window with anti-condensation layer |  |  | ☐ Yes ☐ No |  |
|  | Automatic defrosting system |  |  | ☐ Yes ☐ No |  |
|  | Condensate evaporation system (no external drainage required) |  |  | ☐ Yes ☐ No |  |
|  | Casters: Four swivel casters with brakes for mobility |  |  | ☐ Yes ☐ No |  |
|  | Rapid temperature recovery after door opening |  |  | ☐ Yes ☐ No |  |
|  | Refrigeration System |  |  | ☐ Yes ☐ No |  |
|  | Hermetically sealed, energy-efficient compressor |  |  | ☐ Yes ☐ No |  |
|  | Refrigerant: CFC/HCFC-free, eco-friendly (e.g., R600a or R134a) |  |  | ☐ Yes ☐ No |  |
|  | Alarm and Monitoring System |  |  | ☐ Yes ☐ No |  |
|  | Audible and visual alarms for: |  |  | ☐ Yes ☐ No |  |
|  | High/low temperature |  |  | ☐ Yes ☐ No |  |
|  | Door open >60 seconds |  |  | ☐ Yes ☐ No |  |
|  | Power failure |  |  | ☐ Yes ☐ No |  |
|  | Sensor error |  |  | ☐ Yes ☐ No |  |
|  | Battery backup for alarm and data display (≥8 hours) |  |  | ☐ Yes ☐ No |  |
|  | Alarm test function |  |  | ☐ Yes ☐ No |  |
|  | Remote alarm interface |  |  | ☐ Yes ☐ No |  |
|  | Data Logging and Communication |  |  | ☐ Yes ☐ No |  |
|  | Integrated USB port and chart recorder |  |  | ☐ Yes ☐ No |  |
|  | Data storage: Minimum 30-day temperature history |  |  | ☐ Yes ☐ No |  |
|  | Ethernet or RS-485 port for integration with external monitoring systems |  |  | ☐ Yes ☐ No |  |
|  | Accessories |  |  | ☐ Yes ☐ No |  |
|  | Lockable door with duplicate keys |  |  | ☐ Yes ☐ No |  |
|  | Voltage stabilizer or UPS with minimum 30-minute backup |  |  | ☐ Yes ☐ No |  |
|  | Spare temperature probe and chart paper |  |  | ☐ Yes ☐ No |  |
|  | 2 to 4 stainless steel drawers, easily removable |  |  | ☐ Yes ☐ No |  |
|  | LED internal lighting (automatic ON/OFF) |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
|  | compliant with IEC 60601 or equivalent. |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  |  | ☐ Yes ☐ No |  |
| 27 | Blood Bank Refrigerator |  |  |  |  |
|  | Name of Manufacturer | 1 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Application: Safe storage of whole blood and blood components at +2°C to +6°C, with all accessories |  |  | ☐ Yes ☐ No |  |
|  | Capacity |  |  | ☐ Yes ☐ No |  |
|  | Net internal volume not less than 600 liter |  |  | ☐ Yes ☐ No |  |
|  | Storage capacity: not less than 300–450 blood bags (450 ml each) |  |  | ☐ Yes ☐ No |  |
|  | Modular or adjustable drawer or shelf system for blood bag segregation |  |  | ☐ Yes ☐ No |  |
|  | Blood bag storage baskets or trays made of SS 304 or medical-grade plastic |  |  | ☐ Yes ☐ No |  |
|  | Internal layout |  |  | ☐ Yes ☐ No |  |
|  | 4 to 6 pull-out drawers, adjustable/removable |  |  | ☐ Yes ☐ No |  |
|  | Trays/baskets to hold ≥50–75 blood bags per drawer, depending on configuration |  |  | ☐ Yes ☐ No |  |
|  | Temperature Control |  |  | ☐ Yes ☐ No |  |
|  | Temperature range: +2°C to +6°C, factory set at +4°C |  |  | ☐ Yes ☐ No |  |
|  | Uniform distribution using forced air circulation |  |  | ☐ Yes ☐ No |  |
|  | Accuracy: ±1°C or better |  |  | ☐ Yes ☐ No |  |
|  | Touchscreen control panel 7” with: Real-time temp display (current, min, max) Alarm history, trend graphs, and diagnostic alerts Password-protected access for settings |  |  | ☐ Yes ☐ No |  |
|  | Controller: Microprocessor-based with programmable settings |  |  | ☐ Yes ☐ No |  |
|  | Construction |  |  | ☐ Yes ☐ No |  |
|  | External material: Powder-coated steel or SS 304 |  |  | ☐ Yes ☐ No |  |
|  | Internal chamber: Stainless steel (SS 304) |  |  | ☐ Yes ☐ No |  |
|  | Insulation: CFC-free high-density polyurethane foam (≥75 mm) |  |  | ☐ Yes ☐ No |  |
|  | Door: Double-walled, lockable, with magnetic gasket seal |  |  | ☐ Yes ☐ No |  |
|  | Viewing window: double-glass window with anti-condensation layer |  |  | ☐ Yes ☐ No |  |
|  | Automatic defrosting system |  |  | ☐ Yes ☐ No |  |
|  | Condensate evaporation system (no external drainage required) |  |  | ☐ Yes ☐ No |  |
|  | Fast recovery time after door openings |  |  | ☐ Yes ☐ No |  |
|  | Casters: Four swivel casters with brakes for mobility |  |  | ☐ Yes ☐ No |  |
|  | Refrigeration System |  |  | ☐ Yes ☐ No |  |
|  | Hermetically sealed compressor (low noise, energy efficient) |  |  | ☐ Yes ☐ No |  |
|  | Refrigerant: CFC/HCFC-free, eco-friendly (e.g., R600a or R134a) |  |  | ☐ Yes ☐ No |  |
|  | Alarm & Safety System |  |  | ☐ Yes ☐ No |  |
|  | Visual and audible alarms for: |  |  | ☐ Yes ☐ No |  |
|  | High/low temperature |  |  | ☐ Yes ☐ No |  |
|  | Door open >60 seconds |  |  | ☐ Yes ☐ No |  |
|  | Power failure |  |  | ☐ Yes ☐ No |  |
|  | Sensor error |  |  | ☐ Yes ☐ No |  |
|  | Battery backup for display and alarm (≥8 hours) |  |  | ☐ Yes ☐ No |  |
|  | Alarm test function |  |  | ☐ Yes ☐ No |  |
|  | Remote alarm interface |  |  | ☐ Yes ☐ No |  |
|  | Data Recording & Connectivity |  |  | ☐ Yes ☐ No |  |
|  | Integrated chart recorder and USB data logger |  |  | ☐ Yes ☐ No |  |
|  | Temperature recording for ≥30 days |  |  | ☐ Yes ☐ No |  |
|  | Ethernet or RS485 for central monitoring |  |  | ☐ Yes ☐ No |  |
|  | Accessories |  |  | ☐ Yes ☐ No |  |
|  | Lockable door with duplicate keys |  |  | ☐ Yes ☐ No |  |
|  | Voltage stabilizer or UPS (minimum 30-minute backup) |  |  | ☐ Yes ☐ No |  |
|  | Spare temperature probe and chart paper |  |  | ☐ Yes ☐ No |  |
|  | 4 to 6 stainless steel drawers, easily removable |  |  | ☐ Yes ☐ No |  |
|  | Each drawer with identification slots or labels |  |  | ☐ Yes ☐ No |  |
|  | LED internal lighting (automatic ON/OFF) |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
|  | compliant with IEC 60601 or equivalent. |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  |  | ☐ Yes ☐ No |  |
| 28 | Laboratory Refrigerator |  |  |  |  |
|  | Name of Manufacturer | 3 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Laboratory refrigerator for storage of cultures, media, serums, specimens and general lab storage |  |  | ☐ Yes ☐ No |  |
|  | Single-door, freestanding design, with approximate capacity of 700 Liters |  |  | ☐ Yes ☐ No |  |
|  | Five adjustable shelves |  |  | ☐ Yes ☐ No |  |
|  | Illuminated interior, incandescent lighting |  |  | ☐ Yes ☐ No |  |
|  | Temperature should be variable between 2° C to 8° C or better |  |  | ☐ Yes ☐ No |  |
|  | Interior Finish:Stainless Steel |  |  | ☐ Yes ☐ No |  |
|  | Integrated temperature monitoring system |  |  | ☐ Yes ☐ No |  |
|  | The monitor shall display chamber temperature to the tenth degree centigrade |  |  | ☐ Yes ☐ No |  |
|  | The refrigerator shall have a seven-day temperature chart recorder. |  |  | ☐ Yes ☐ No |  |
|  | The laboratory refrigerator shall have a security locking system |  |  | ☐ Yes ☐ No |  |
|  | ALARM |  |  | ☐ Yes ☐ No |  |
|  | The refrigerator should have an adjustable high/low temperature alarm |  |  | ☐ Yes ☐ No |  |
|  | The refrigerator should have an audible and visual alarm to alert staff to power loss and temperature deviation |  |  | ☐ Yes ☐ No |  |
|  | The refrigerator will have a method to manually test alarms. |  |  | ☐ Yes ☐ No |  |
|  | Audible alarm should not be defeatable, but rather temporarily silenced. A corresponding visual indicator should be ON during temporary alarm silencing period. |  |  | ☐ Yes ☐ No |  |
|  | The refrigerator shall have provision for contacts to interface with the Building Management System (BMS) |  |  | ☐ Yes ☐ No |  |
|  | Microprocessor with Backup Battery |  |  | ☐ Yes ☐ No |  |
|  | To incorporate four casters, at least two of which should be lockable |  |  | ☐ Yes ☐ No |  |
|  | The refrigerator shall have a magnetic door gasket for positive seal |  |  | ☐ Yes ☐ No |  |
|  | Refrigeration System |  |  | ☐ Yes ☐ No |  |
|  | Medical grade Hermetically sealed, air-cooled compressor |  |  | ☐ Yes ☐ No |  |
|  | Non-CFC refrigerant |  |  | ☐ Yes ☐ No |  |
|  | Forced air circulation maintains chamber uniformity of +/-1°C and provides quick recovery. |  |  | ☐ Yes ☐ No |  |
|  | Interior fans shut down when door is opened. |  |  | ☐ Yes ☐ No |  |
|  | Automatic condensate evaporation system |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  |  | ☐ Yes ☐ No |  |
| 29 | Blood Bag Refrigerated Centrifuge |  |  |  |  |
|  | Name of Manufacturer | 2 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Capacity and Rotors |  |  | ☐ Yes ☐ No |  |
|  | Swing-out rotor with capacity for:At least 12–16 whole blood bags (450 ml) or equivalent buckets |  |  | ☐ Yes ☐ No |  |
|  | Compatible with single, double, or triple bag systems |  |  | ☐ Yes ☐ No |  |
|  | Adapters/inserts for 250 ml, 350 ml, 450 ml, and 500 ml bags |  |  | ☐ Yes ☐ No |  |
|  | Rotor and buckets must be autoclavable or easy to disinfect |  |  | ☐ Yes ☐ No |  |
|  | Buckets must be sealed or capped to prevent aerosol contamination |  |  | ☐ Yes ☐ No |  |
|  | Speed and RCF |  |  | ☐ Yes ☐ No |  |
|  | Maximum speed: At least 5,000 rpm |  |  | ☐ Yes ☐ No |  |
|  | Maximum RCF (Relative Centrifugal Force): Minimum 5,000 × g |  |  | ☐ Yes ☐ No |  |
|  | Speed increment: Adjustable in ≥10 rpm steps |  |  | ☐ Yes ☐ No |  |
|  | Programmable acceleration and deceleration: At least 10 profiles |  |  | ☐ Yes ☐ No |  |
|  | Soft spin modes to preserve blood cell integrity |  |  | ☐ Yes ☐ No |  |
|  | Temperature Control |  |  | ☐ Yes ☐ No |  |
|  | Refrigerated system with preset operating temperature of +4°C |  |  | ☐ Yes ☐ No |  |
|  | Adjustable temperature range: -5°C to +25°C |  |  | ☐ Yes ☐ No |  |
|  | Temperature accuracy: ±1°C |  |  | ☐ Yes ☐ No |  |
|  | CFC/HCFC-free refrigerant (e.g., R600a or R404a) |  |  | ☐ Yes ☐ No |  |
|  | Automatic defrosting and silent compressor operation |  |  | ☐ Yes ☐ No |  |
|  | Control System |  |  | ☐ Yes ☐ No |  |
|  | Touchscreen (7” or larger) or advanced digital control panel |  |  | ☐ Yes ☐ No |  |
|  | Display of: Real-time speed (RPM) and RCF,Set and actual temperature, Time countdown, Rotor status and cycle progress |  |  | ☐ Yes ☐ No |  |
|  | Programmable memory for ≥10 user-defined protocols |  |  | ☐ Yes ☐ No |  |
|  | Timer range: 1–99 minutes or continuous run |  |  | ☐ Yes ☐ No |  |
|  | Safety Features |  |  | ☐ Yes ☐ No |  |
|  | Lid locking system: Lid cannot be opened while rotor is in motion |  |  | ☐ Yes ☐ No |  |
|  | Automatic lid lock and unlock controlled electronically |  |  | ☐ Yes ☐ No |  |
|  | Manual/emergency lid release in case of power failure |  |  | ☐ Yes ☐ No |  |
|  | Imbalance detection with auto shut-off |  |  | ☐ Yes ☐ No |  |
|  | Over-speed, over-temperature, and door open protection |  |  | ☐ Yes ☐ No |  |
|  | Rotor auto-recognition system |  |  | ☐ Yes ☐ No |  |
|  | Lid inspection window made from impact-resistant transparent polycarbonate or glass, anti-fog treated, if manufacturer supports it |  |  | ☐ Yes ☐ No |  |
|  | Construction and Design |  |  | ☐ Yes ☐ No |  |
|  | Outer casing: Powder-coated steel or stainless steel |  |  | ☐ Yes ☐ No |  |
|  | Inner chamber: SS 304 grade stainless steel, corrosion-resistant |  |  | ☐ Yes ☐ No |  |
|  | Rotor: High-strength aluminum or fiber composite |  |  | ☐ Yes ☐ No |  |
|  | Vibration reduction: Heavy-duty anti-vibration system |  |  | ☐ Yes ☐ No |  |
|  | Built-in shock absorbers and low-noise operation |  |  | ☐ Yes ☐ No |  |
|  | Data and Connectivity |  |  | ☐ Yes ☐ No |  |
|  | Built-in cycle logging for last 50–100 runs |  |  | ☐ Yes ☐ No |  |
|  | Accessories |  |  | ☐ Yes ☐ No |  |
|  | One complete swing-out rotor with 4–8 sealed buckets |  |  | ☐ Yes ☐ No |  |
|  | adapters for various blood bag sizes |  |  | ☐ Yes ☐ No |  |
|  | Voltage stabilizer or UPS (at least 30 minutes backup) |  |  | ☐ Yes ☐ No |  |
|  | Spare rotor seal and calibration certificate |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  |  | ☐ Yes ☐ No |  |
|  | compliant with IEC 60601 or equivalent. |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  |  | ☐ Yes ☐ No |  |
| 30 | Platelet Incubators, Benchtop |  |  |  |  |
| 1 | Name of Manufacturer | 1 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below1 |
| 2 | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
| 3 | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
| 4 | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
| 5 | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
| 6 | TEMPERATURE CONTROL WITH UNIFORM AIR CIRCULATION |  |  | ☐ Yes ☐ No |  |
| 7 | Designed to be used in conjunction with agitator |  |  | ☐ Yes ☐ No |  |
| 8 | Digital speed selection and timer |  |  | ☐ Yes ☐ No |  |
| 9 | Speed RPM:40-80 Approx. |  |  | ☐ Yes ☐ No |  |
| 10 | Time: 0 to 99 min |  |  | ☐ Yes ☐ No |  |
| 11 | Touch screen alarm and control system |  |  | ☐ Yes ☐ No |  |
| 12 | 7-day chart recorder |  |  | ☐ Yes ☐ No |  |
| 13 | Dual temperature display |  |  | ☐ Yes ☐ No |  |
| 14 | AUDIO VISUAL ALERT ALARMS |  |  | ☐ Yes ☐ No |  |
| 15 | Motion sensor |  |  | ☐ Yes ☐ No |  |
| 16 | Temperature alarm |  |  | ☐ Yes ☐ No |  |
| 17 | Door open alarm |  |  | ☐ Yes ☐ No |  |
| 18 | Remote alarm contacts |  |  | ☐ Yes ☐ No |  |
| 19 | Access port |  |  | ☐ Yes ☐ No |  |
| 20 | Capacity : 50 bags |  |  | ☐ Yes ☐ No |  |
| 21 | SS internal chamber and internal shelves |  |  | ☐ Yes ☐ No |  |
| 22 | Glass transparent door, with lock |  |  | ☐ Yes ☐ No |  |
| 23 | Online UPS |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
| 24 | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  |  | ☐ Yes ☐ No |  |
| 25 | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  |  | ☐ Yes ☐ No |  |
| 31 | Blood Warmer |  |  |  |  |
|  | Name of Manufacturer | 4 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Microprocessor controlled dry heat blood transfusion warmer |  |  | ☐ Yes ☐ No |  |
|  | TEMP. SETTING degree C 38 - 42 degree C, accuracy (+/- 0.5 deg. C) and resolution (0.5 deg. C or better) |  |  | ☐ Yes ☐ No |  |
|  | Variable flow rate ≥ 70 ml / min. Flow rate shall be maintained (continuous, constant and stable). |  |  | ☐ Yes ☐ No |  |
|  | Output temperature rise from approx. 10 deg. C to not less than 35 deg. C shall be maintained constant and stable throughout the flow range. |  |  | ☐ Yes ☐ No |  |
|  | Warm-up time: less than 60 sec. |  |  | ☐ Yes ☐ No |  |
|  | Digital Display parameters: |  |  | ☐ Yes ☐ No |  |
|  | Output temperature |  |  | ☐ Yes ☐ No |  |
|  | Flow rate in ml / min |  |  | ☐ Yes ☐ No |  |
|  | Error messages |  |  | ☐ Yes ☐ No |  |
|  | Heating indicator |  |  | ☐ Yes ☐ No |  |
|  | Heater power,Specify |  |  | ☐ Yes ☐ No |  |
|  | Specify high temperature safety cutoff |  |  | ☐ Yes ☐ No |  |
|  | Audio visual alarms and indicators: |  |  | ☐ Yes ☐ No |  |
|  | Temperature variations (high and low) |  |  | ☐ Yes ☐ No |  |
|  | System functional errors or internal failures |  |  | ☐ Yes ☐ No |  |
|  | Empty cartridge (end of transfusion) |  |  | ☐ Yes ☐ No |  |
|  | Low flow |  |  | ☐ Yes ☐ No |  |
|  | Air in line |  |  | ☐ Yes ☐ No |  |
|  | Safety features : |  |  | ☐ Yes ☐ No |  |
|  | Splash proof panel |  |  | ☐ Yes ☐ No |  |
|  | Internal self-test |  |  | ☐ Yes ☐ No |  |
|  | Infusion halt on alarm |  |  | ☐ Yes ☐ No |  |
|  | IV pole attachment or clamp should be incorporated |  |  | ☐ Yes ☐ No |  |
|  | Systems that operate using standard IV tubing and blood administration sets are preferable. However, if a warming set is required, then any disposable items necessary for the operation of the unit should be mentioned and listed and priced seperately |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  |  | ☐ Yes ☐ No |  |
| 32 | Warmer, Fluid/ Blood, Portable |  |  |  |  |
|  | Name of Manufacturer | 2 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Microprocessor controlled dry heat blood transfusion warmer |  |  | ☐ Yes ☐ No |  |
|  | The unit should have user adjustable temperature setting. Specify range (or value if fixed set point), accuracy (+/- 0.5 deg. C) and resolution (0.5 deg. C or better) |  |  | ☐ Yes ☐ No |  |
|  | Variable flow rate up to ≥ 100 ml / min. Flow rate shall be maintained (continuous, constant and stable). |  |  | ☐ Yes ☐ No |  |
|  | Output temperature rise from approx. 10 deg. C to not less than 35 deg. C shall be maintained constant and stable throughout the flow range. Specify detailed data pertaining to your system. |  |  | ☐ Yes ☐ No |  |
|  | The unit shall digitally display important operating parameters such as: |  |  | ☐ Yes ☐ No |  |
|  | Output temperature (specify resolution and accuracy) |  |  | ☐ Yes ☐ No |  |
|  | Flow rate in ml / min (specify resolution and accuracy) |  |  | ☐ Yes ☐ No |  |
|  | Error messages (list all available messages) |  |  | ☐ Yes ☐ No |  |
|  | Heating indicator |  |  | ☐ Yes ☐ No |  |
|  | Specify heater power |  |  | ☐ Yes ☐ No |  |
|  | Specify high temperature safety cutoff and breaking method |  |  | ☐ Yes ☐ No |  |
|  | System safety features should include audio visual alarms and indicators for: |  |  | ☐ Yes ☐ No |  |
|  | Temperature variations (high and low) |  |  | ☐ Yes ☐ No |  |
|  | System functional errors or internal failures |  |  | ☐ Yes ☐ No |  |
|  | Empty cartridge (end of transfusion) |  |  | ☐ Yes ☐ No |  |
|  | Low flow |  |  | ☐ Yes ☐ No |  |
|  | Air in line |  |  | ☐ Yes ☐ No |  |
|  | Safety features should also include: |  |  | ☐ Yes ☐ No |  |
|  | Splash proof panel |  |  | ☐ Yes ☐ No |  |
|  | Internal self-test |  |  | ☐ Yes ☐ No |  |
|  | Infusion halt on alarm |  |  | ☐ Yes ☐ No |  |
|  | Maintenance-free operation |  |  | ☐ Yes ☐ No |  |
|  | Compliance with safety standards such as IEC 601-1 or similar equivalent (specify) |  |  | ☐ Yes ☐ No |  |
|  | IV pole attachment or clamp should be incorporated |  |  | ☐ Yes ☐ No |  |
|  | Warm-up time: less than 60 sec. |  |  | ☐ Yes ☐ No |  |
|  | Systems that operate using standard IV tubing and blood administration sets are preferable. However, if a warming set is required, then any disposable items necessary for the operation of the unit should be mentioned and listed |  |  | ☐ Yes ☐ No |  |
| 33 | Warmer, Infant, Care System |  |  |  |  |
|  | Name of Manufacturer | 1 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Microprocessor controlled heater |  |  | ☐ Yes ☐ No |  |
|  | Even heat distribution over the bed section |  |  | ☐ Yes ☐ No |  |
|  | Manual mode (multi-level) |  |  | ☐ Yes ☐ No |  |
|  | Baby skin mode (Servo), adjustable set temperature: 34 – 38 °C |  |  | ☐ Yes ☐ No |  |
|  | Swivel mechanism in either direction to facilitate access to baby, x-ray imaging…etc |  |  | ☐ Yes ☐ No |  |
|  | clear text message display |  |  | ☐ Yes ☐ No |  |
|  | integrated examination light |  |  | ☐ Yes ☐ No |  |
|  | alarming (for parameters out-of-limit) and self-test system |  |  | ☐ Yes ☐ No |  |
|  | bed section Tiltable from +10° to -10° or better with foldable sidewalls |  |  | ☐ Yes ☐ No |  |
|  | Electrically foot controlled height adjustable |  |  | ☐ Yes ☐ No |  |
|  | With storage drawers |  |  | ☐ Yes ☐ No |  |
|  | Mobile with 4 swiveling castors (2 with brakes) |  |  | ☐ Yes ☐ No |  |
|  | Accessories, the following shall be included: |  |  | ☐ Yes ☐ No |  |
|  | Skin temperature sensor (Qty: 2) |  |  | ☐ Yes ☐ No |  |
|  | Adhesive temperature pads (Qty: 20) |  |  | ☐ Yes ☐ No |  |
|  | Bed Mattress |  |  | ☐ Yes ☐ No |  |
|  | Shall be CE marked and/or FDA approved |  |  | ☐ Yes ☐ No |  |
| 34 | Blender, Gas, Air/Oxygen |  |  |  |  |
|  | Name of Manufacturer | 1 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | The regulator shall be impact-resistant, durable and constructed of high quality materials |  |  | ☐ Yes ☐ No |  |
|  | Inlet supply pressure (for Air and Oxygen): 3-4 bar |  |  | ☐ Yes ☐ No |  |
|  | Connected to central medical air system and oxygen outlet via hoses with British standard probe to BS 5682:2015 (unless otherwise requested by the client) |  |  | ☐ Yes ☐ No |  |
|  | O2 concentration control: 21-100% |  |  | ☐ Yes ☐ No |  |
|  | Blender flow rate control: 0-10 l/min |  |  | ☐ Yes ☐ No |  |
|  | With audible alarm in case of high inlet pressure or absence of either gas supply |  |  | ☐ Yes ☐ No |  |
|  | All surfaces shall be disinfectable |  |  | ☐ Yes ☐ No |  |
|  | All required accessories, consumables, hoses/tubes etc for full functionality shall be included |  |  | ☐ Yes ☐ No |  |
|  | The unit shall be CE marked and/or FDA approved |  |  | ☐ Yes ☐ No |  |
| 35 | Medical Refrigerator |  |  |  |  |
|  | Manufacturer | 2 |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting from the date of installation and Final Acceptance |  |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Douple-door, freestanding design, with approximate capacity of 700 Liters |  |  | ☐ Yes ☐ No |  |
|  | Five adjustable shelves |  |  | ☐ Yes ☐ No |  |
|  | Illuminated interior, LED lighting |  |  | ☐ Yes ☐ No |  |
|  | Temperature should be variable between 2° C to 8° C or better |  |  | ☐ Yes ☐ No |  |
|  | Integrated temperature monitoring system |  |  | ☐ Yes ☐ No |  |
|  | The monitor shall display chamber temperature to the tenth degree centigrade |  |  | ☐ Yes ☐ No |  |
|  | The refrigerator shall have a seven-day temperature chart recorder. |  |  | ☐ Yes ☐ No |  |
|  | The refrigerator shall have a security locking system |  |  | ☐ Yes ☐ No |  |
|  | The refrigerator should have an adjustable high/low temperature alarm |  |  | ☐ Yes ☐ No |  |
|  | 8. The refrigerator should have an audible and visual alarm to alert staff to power loss and temperature deviation |  |  | ☐ Yes ☐ No |  |
|  | The refrigerator will have a method to manually test alarms. |  |  | ☐ Yes ☐ No |  |
|  | Audible alarm should not be defeatable, but rather temporarily silenced. A corresponding visual indicator should be ON during temporary alarm silencing period. |  |  | ☐ Yes ☐ No |  |
|  | The refrigerator shall have provision for contacts to interface with the Building Management System (BMS) |  |  | ☐ Yes ☐ No |  |
|  | To incorporate four casters, at least two of which should be lockable |  |  | ☐ Yes ☐ No |  |
|  | The refrigerator shall have a magnetic door gasket for positive seal |  |  | ☐ Yes ☐ No |  |
|  | Refrigeration System |  |  | ☐ Yes ☐ No |  |
|  | Hermetically sealed, air-cooled compressor |  |  | ☐ Yes ☐ No |  |
|  | Non-CFC refrigerant |  |  | ☐ Yes ☐ No |  |
|  | Forced air circulation maintains chamber uniformity of +/-1°C and provides quick recovery. |  |  | ☐ Yes ☐ No |  |
|  | Interior fans shut down when door is opened. |  |  | ☐ Yes ☐ No |  |
|  | Automatic condensate evaporation system |  |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  |  | ☐ Yes ☐ No |  |

Delivery requirements –– Comparative Data Table

| UNOPS Requirements |  | Is quotation compliant? Bidder to complete | Details Bidder to complete |
| --- | --- | --- | --- |
| Delivery time | Bidder shall deliver the goods [120 calendar days] after Contract signature. | ☐ Yes ☐ No | Insert details |
| Delivery place and Incoterms rules | Delivery at Place DAP to Hospitals and Institutes mentioned in Bills of quantity tables- Form B, in case that the site is not ready, the supplier agrees to store the equipment in their warehouse until the site is prepared to receive the equipment. Incoterms rules as per Incoterms 2020. | ☐ Yes ☐ No | Insert details |
| Consignee details | This Shall include the Installation and Training requirements. Consignee will be the UNOPS and MoH representative | ☐ Yes ☐ No | Insert details |
| UNOPS Right to vary requirements | At the time the Contract is awarded, UNOPS reserves the right to vary the quantity of the goods and associated services specified above, provided this does not exceed +/- [20%] , without any change in the unit prices or other terms and conditions of the RFQ. | ☐ Yes ☐ No | Insert details |

00ax: +45 45 33 75 01

The offered goods and related services (if applicable) are in accordance with the required specifications and requirements specified in Section II: Schedule of Requirements.

☐ Yes   ☐ No

ANY DEVIATION MUST BE LISTED BELOW:

_________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

Name	: _____________________________________________________________

Title	: _____________________________________________________________

Date	: _____________________________________________________________

Signature	: _____________________________________________________________

If subcontracting is anticipated in the performance by the contractor, Procurement Official should insert the following:

List of subcontractors or suppliers

Bidder must identify the names of all subcontractors/suppliers who will be providing good/services under this contract and the type of work being subcontracted, if applicable.

- _[Full legal name and address of subcontractors]___________

- _________________________________________________

(C)_________________________________________________

I, the undersigned, certify that I am duly authorized by [insert full name of bidder] to sign this bid and bind [insert full name of bidder] should UNOPS accept this bid:

Name	: _____________________________________________________________

Title	: _____________________________________________________________

Date	: _____________________________________________________________

Signature	: _____________________________________________________________

# Form D: Previous Experience Form

RFQ reference no: [Insert UNOPS tender reference number]

Name of Bidder: [insert name of Bidder]

| Description of services/goods | Country | Total amount of Contract | Contract Identification and Title and Contact details of Client (Name, Address, telephone, email, fax) | Year project was undertaken |
| --- | --- | --- | --- | --- |

Name	: _____________________________________________________________

Title	: _____________________________________________________________

Date	: _____________________________________________________________

Signature	: _____________________________________________________________

# Form E: Self Disclosure Form

RFQ reference no: [insert RFQ reference No.]

Name of Bidder: [insert name of Bidder]

| Ref. | Critical area | Response | Comments |
| --- | --- | --- | --- |
| 1 | Has the entity or individual declared bankruptcy, or been involved in bankruptcy or receivership proceedings, or is there any judgment or pending legal action against them, which could impair operations in the foreseeable future; | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 2 | Has the entity or individual been found (or is awaiting a judgment or administrative decision) in breach of their obligations relating to the payment of taxes or social security contributions? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 3 | Has the entity or individual been found (or is awaiting a judgment or administrative decision) guilty of misconduct for the violation of applicable laws, regulations or ethical standards of the profession to which they belong? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 4 | Has the entity or individual engaged, or attempted to engage, in any Proscribed Practices in the past? For the purposes of this provision, “Proscribed Practices” are defined in the UNOPS Operational Instruction on Vendor Sanctions and include the following: - A corrupt practice: the offering, giving, receiving, or soliciting, directly or indirectly, anything of value to influence improperly the actions of another party; - A fraudulent practice: any act or omission, including a misrepresentation, that knowingly or recklessly misleads, or attempts to mislead, a party to obtain a financial or other benefit or to avoid an obligation; - A coercive practice: an act or omission that impairs or harms, or threatens to impair or harm, directly or indirectly, any party or the property of the party to improperly influence the actions of a party; - A collusive practice: an arrangement between two or more parties designed to achieve an improper purpose, including influencing improperly the actions of another party; - An unethical practice: Conduct or behaviour that is contrary to the conflict of interest, gifts and hospitality, post-employment provisions or other published requirements of doing business with UNOPS; and - Obstruction: Acts or omissions by a Vendor that prevent or hinder UNOPS from investigating instances of possible Proscribed Practices. | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 5 | Has the entity or individual unduly obtained, or attempted to unduly obtain, any confidential information in connection with this procurement process and any agreement that may be awarded as a result of this process? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 6 | Does the entity or individual have a conflict of interest, as provided in the Instructions to bidders, Article 3 “Bidder eligibility”, that may prevent them from entering into an agreement with UNOPS, or that may in any way jeopardize their performance in carrying out their obligations under the agreement? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 7 | Has the entity or individual been found (or is awaiting a judgment or administrative decision) to be involved in any of the following: - Fraudulent practice; - Corrupt practice; - Affiliation with a criminal organization; - Money laundering; - Terrorist financing; - Child labour; or - Human trafficking? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 8 | Has the entity or individual had significant performance issues in the past that led to the early termination of a legal commitment or the application of damages or other contractual penalties, or that were discovered following inspections, audits or investigations? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 9 | Has the entity or individual committed any financial irregularity, resulting from an act or omission or the inclusion of an unjustified item of expenditure outside a contractual engagement that had the effect of prejudicing their use of public funds? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 10 | Has the entity or individual been found (or is awaiting a judgment or administrative decision) to have created an entity in a different jurisdiction with the intent of circumventing fiscal, social or any other legal obligations in the jurisdiction of their registered office, central administration or principal place of business? | Select | [If the response is ‘Yes’, provide an explanation here.] |

I, the undersigned, hereby declare that the information provided is true and correct. I also understand that any willful dishonesty may result in the refusal of this submission. I understand that engagement in any of the above critical areas may, at the discretion of UNOPS, automatically result in the exclusion from this procurement process.

| Name: |  |
| --- | --- |
| Title |  |
| Date: | Sep 10, 2024 |
| Signature |  |

[Stamp this form with the official stamp of the bidder]

Form F: Joint Venture Partner Information Form

[The Bidder shall fill in this Form in accordance with the instructions indicated below].

RFQ reference no: [insert RFQ reference No.]

Name of Bidder: [insert name of bidder]

Date: [insert submission date]

To be completed and returned with your Bid if the Bid is submitted as a Joint Venture/Consortium/Association.

| JV / Consortium/ Association Information |  |
| --- | --- |
| Name | [complete] |
| Names of each partner and contact information (address, telephone numbers, fax numbers, e-mail address) | [complete] |
| Name of leading partner (with authority to bind the JV, Consortium, Association during the Bidding process and, in the event a Contract is awarded, during contract execution) | [complete] |
| Proposed proportion of responsibilities between partners (in %) with indication of the type of the goods/services to be delivered by each | [complete] |

Signatures of all partners of the JV:

We hereby confirm that if the contract is awarded, all parties of the Joint Venture/Consortium/Association shall be jointly and severally liable to UNOPS for the fulfillment of the provisions of the Contract.

Name of partner: ________________________		Name of partner: _________________________

Signature: ______________________________		Signature: _______________________________

Date: _______________________				Date: ________________________

Name of partner: ________________________		Name of partner: _________________________

Signature: ______________________________		Signature: _______________________________

Date: _______________________				Date: ________________________

Form G: Manufacturer’s Authorization Form

A letter issued by the manufacturer authorizing the applicant to participate in this particular RFQ must be submitted with the bid in the format provided in this Form.

To be eligible for delivery of goods, the bidder must be either the manufacturer of the offered goods or a sole representative of the manufacturer to the United Nations. Should offers for a particular make and model be received from more than one appointed representative, UNOPS reserves the right to select only one.

RFQ reference no: [Insert UNOPS tender reference number]

Name of Bidder: [insert name of bidder]

Date: [insert submission date]

To: [bidder to insert]

WHEREAS

We [insert complete name of manufacturer], who are official manufacturers of [insert type of goods manufactured], having factories at [insert full address of manufacturer’s factories], do hereby authorize [insert complete name of bidder] to submit a bid the purpose of which is to provide the following goods, manufactured by us [insert name and or brief description of the goods], and to subsequently negotiate and sign the contract.

We hereby extend our full guarantee and warranty in accordance with Clause 13 of the General Conditions for Goods, with respect to the goods offered by the above firm.

Signed: [insert signature(s) of authorized representative(s) of the manufacturer]

Name: [insert complete name(s) of authorized representative(s) of the manufacturer]

Title: [insert title]

Dated on ____________ day of __________________, _______ [insert date of signing]