Section III: Returnable Bidding Forms

eSourcing reference:   RFQ/2026/62289]

Note to Bidders: The following returnable forms are part of this RFQ and must be completed and returned by bidders as part of their quotation. Instructions to complete each Form are highlighted in blue in each Form. Please complete the Returnable Bidding Forms as instructed and return them as part of your quotation by uploading them against their specific Document Checklist in the UNOPS eSourcing system.

This Section comprises the following Returnable Bidding Forms:

- Form A: Quotation Submission Form

- Form B: Price Schedule Form

- Form C: Technical Quotation Form

- Form D: Previous Experience Form

- Form E: Self Disclosure Form

- Form F: Joint Venture Partner Information Form

- Form G: Manufacturer’s authorization form

Form A: Quotation submission form

Bidders are requested to complete this form, sign it and return it as part of their bid submission. The bidder shall fill in this form in accordance with the instructions indicated. No alterations to its format shall be permitted and no substitutions shall be accepted.

Date: [Insert submission date]

Subject: Quotation for the supply of [Insert a brief description of goods/services] in [Name of country/city], RFQ Case No. [Insert RFQ ref number], dated [insert date]

We, the undersigned, declare that:

  - We offer to supply in conformity with the bidding documents, including the UNOPS General Conditions of Contract;

  - Our quotation shall be valid for the period of time of [insert number of days which shall not be less than the specified in the Tender Particulars section, Period of Validity of Quotations] from the date fixed for the submission deadline as set out in the RFQ, and it shall remain binding upon us and may be accepted at any time before the expiration of that period;

  - We have no conflict of interest in any activity that would put it, if selected for this assignment, in a conflict of interest with UNOPS [If you have any actual or potential conflict of interest as defined in Article 3 of Section I: Instructions to Bidders, please disclose it here];;

  - Our firm confirms that the offeror and sub-contractors have not been associated, or had been involved in any way, directly or indirectly, with the preparation of the design, terms of references and/or other documents used as a part of this solicitation;

  - Our firm, its affiliates or subsidiaries—including any subcontractors or suppliers for any part of the Contract—has not been declared ineligible by UNOPS, nor is included in the suspended/ineligibility list of the UN/PD, other UN Agencies, the UN Security Council, and the World Bank, in accordance with Instructions to Bidders Article 3, Eligibility;

  - We embrace the UN Supplier Code of Conduct and adhere to the principles of the UN Global Compact;

  - We have not declared bankruptcy, are not involved in bankruptcy or receivership proceedings, and there is no judgment or pending legal action against them that could impair their operations in the foreseeable future;

  - We have not offered and will not offer fees, gifts and/or favours of kind in exchange for this RFQ and will not engage in any such activity during the performance of any Contract awarded.

I, the undersigned, certify that I am duly authorized by [insert full name of bidder] to sign this quotation and bind [insert full name of bidder] should UNOPS accept this quotation:

Name: [complete]

Title: [complete]

Date: [complete]

Signature: _____________________________________________________________

Provide the name and contact information for the primary contact from your company for this quotation:

Name: [complete]

Title: [complete]

Email address: [complete]

Telephone: [complete]

# Form B: Price Schedule Form

Bidders shall fill in this Price Schedule Form in accordance with the instructions indicated.

**Please fill in all your prices in the attached excel sheet and send it as (excel book) along with the below table**

Pricing Methodology:

- In addition to DAP price bidder to add a breakdown of prices: cost of goods cleared for export (FCA) separate from transport and insurance.

- The Supplier shall, within fourteen (14) days of Contract signature, provide a contingency plan for alternative routing (e.g., land corridors or bypass ports) should the primary maritime route be restricted.

- To confirm whether the primary components and finished goods are sourced from regions not currently subject to maritime blockades or significant transit disruptions in the Strait of Hormuz or Red Sea (Geographic Origin Declaration).

RFQ reference no: [insert RFQ reference No.]

| Currency | USD |
| --- | --- |

Bills of quantity Lot# 5: Supply & installation and commissioning ICUs medical equipment

| Item No | Description |  | Qty | Locations and Distributions | Unit Price DAP | Total Price DAP |
| --- | --- | --- | --- | --- | --- | --- |
|  | Lung Ventilator For Adult & Pediatric |  | 5 | Ataq TFC, Shabwah: 2 Al-Mansoura TFC, Aden:2 Hais TFC, Al Hudayda: 1 | insert | insert |
|  | Bedhead unit Horizontal, ICU |  | 2 | Ataq TFC, Shabwah:2 | insert | insert |
|  | Bedhead unit Horizontal, Patient Room |  | 2 | Ataq TFC, Shabwah:2 |  |  |
|  | Bedhead unit Vertical, Patient Room |  | 2 | Ataq TFC, Shabwah:2 |  |  |
|  | Defibrillator Biphasic Adult/Pediatric |  | 2 | Ataq TFC, Shabwah:1 Al-Mansoura TFC, Aden:1 | insert | insert |
|  | Patient Monitoring System (Bedside Monitors + Central Station) |  | 4 | Ataq TFC, Shabwah:2 Al-Mansoura TFC, Aden:2 | insert | insert |
|  | Electrocardiography, (ECG) |  | 4 | Ataq TFC, Shabwah:1 Al-Mansoura TFC, Aden:1 Al Hezma, Mariab:1 Radfan Hospital - Lahj:1 | insert | insert |
|  | Ultrasonic Nebulizer with Trolly |  | 4 | Ataq TFC, Shabwah:2 Al-Mansoura TFC, Aden:2 | insert | insert |
|  | Electrical ICU Bed 4 |  | 10 | Ataq TFC, Shabwah:2 Al-Mansoura TFC, Aden:2 Al Hezma, Mariab: 6 | insert | insert |
|  | Mobilel Suction Unit 2*2.5L Suitable for ICU |  | 4 | Ataq TFC, Shabwah:2 Al-Mansoura TFC, Aden:2 | insert | insert |
|  | Bed, Patient, Hydraulic |  | 21 | Ataq TFC, Shabwah: 4 Al-Mansoura TFC, Aden:4 Hais TFC, Al Hudayda: 3 Al Hezma, Mariab:10 | insert | insert |
|  | Pulse Oxmeter (Handheld) |  | 1 | Al Hezma, Mariab: 1 | insert | insert |
|  | Medical Oxygen Concentrator |  | 4 | Ataq TFC, Shabwah:2 Al-Mansoura TFC, Aden:2 | insert | insert |
|  | Resuscitation Set (Adult & Child) |  | 4 | Ataq TFC, Shabwah:2 Al-Mansoura TFC, Aden:2 | insert | insert |
|  | Electrosurgical Unit, Bipolar |  | 5 | Ibn Sina Hospital, Al Mukalla:4 Al Hezma, Mariab: 1 | insert | insert |
|  | Monitor, Physiologic, Anesthesia |  | 4 | Ibn Sina Hospital, Al Mukalla | insert | insert |
|  | Table, Surgical, Major |  | 4 | Ibn Sina Hospital, Al Mukalla | insert | insert |
|  | Monitor, Physiologic, Portable |  | 4 | Al Hezma, Mariab: 4 | insert | insert |
|  | Bassinet, LDRP |  | 1 | Al Hezma, Mariab | insert | insert |
|  | Bed, Birthing |  | 1 | Al Hezma, Mariab | insert | insert |
|  | Light, Exam/Procedure, Single, Ceiling |  | 1 | Al Hezma, Mariab | insert | insert |
|  | Light, Phototherapy, Infant, Mobile |  | 1 | Al Hezma, Mariab | insert | insert |
|  | Monitor, O.B., Intrapartum, Maternal/Fetal |  | 1 | Al Hezma, Mariab | insert | insert |
|  | Patient Monitor, Physiologic |  | 1 | Al Hezma, Mariab | insert | insert |
|  | Regulator, Suction, Intermittent/Continuous |  | 4 | Al Hezma, Mariab | insert | insert |
|  | Table, Surgical, Minor |  | 1 | Al Hezma, Mariab | insert | insert |
|  | Anesthesia Machine Minor with monitor |  | 1 | Al Hezma, Mariab | insert | insert |
|  |  | Total Price |  |  |  | insert |

Payment terms 30 days accepted: ☐ Yes

Bidder’s discount for accelerated payment: ____% of total firm price for each calendar day less than thirty (30) days

List of subcontractors or suppliers

Bidder must identify the names of all subcontractors/suppliers who will be providing good/services under this Contract and the type of work being subcontracted, if applicable.

- _[Full legal name and address of subcontractors]___________

- _________________________________________________

- _________________________________________________

I, the undersigned, certify that I am duly authorized by [insert full name of Bidder] to sign this quotation and bind [insert full name of Bidder] should UNOPS accept this quotation:

Name	: _____________________________________________________________

Title	: _____________________________________________________________

Date	: _____________________________________________________________

Signature	: _____________________________________________________________

| Item No | UNOPS minimum technical requirements | Qty | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete |
| --- | --- | --- | --- | --- |
| 1 | Ventilator Machine |  |  |  |
|  | Manufacturer | 5 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Country where the manufacturer is based |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting from the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Ventilator, Intensive Care |  | ☐ Yes ☐ No |  |
|  | AC and battery powered |  | ☐ Yes ☐ No |  |
|  | Backup battery |  | ☐ Yes ☐ No |  |
|  | Colored Touch Screen not less than 12 |  | ☐ Yes ☐ No |  |
|  | Central Gas inlets (O2 & Air) |  | ☐ Yes ☐ No |  |
|  | Each high-pressure input port with a filter and water trap. |  | ☐ Yes ☐ No |  |
|  | The machine can be operated by both central gas supply and gas cylinders |  | ☐ Yes ☐ No |  |
|  | Electronic nebulizer |  | ☐ Yes ☐ No |  |
|  | Internal function testing/leak testing. |  | ☐ Yes ☐ No |  |
|  | Event log for errors traceability |  | ☐ Yes ☐ No |  |
|  | O2 cell |  | ☐ Yes ☐ No |  |
|  | Active Expiratory Valve (Autoclavable) |  | ☐ Yes ☐ No |  |
|  | Compatible active humidifying system with servo-controlled heating and open system with two Mods: One for NIV and the another for Invasive ventilation. |  | ☐ Yes ☐ No |  |
|  | At least IP21 degree of protection to the harmful ingress of water (fluid spill resistance). |  | ☐ Yes ☐ No |  |
|  | Ventilation modes |  | ☐ Yes ☐ No |  |
|  | Pressure control ventilation (PCV). |  | ☐ Yes ☐ No |  |
|  | Volume control ventilation (VCV). |  | ☐ Yes ☐ No |  |
|  | Pressure support ventilation (PSV). |  | ☐ Yes ☐ No |  |
|  | Synchronized intermittent mandatory ventilation (SIMV) . |  | ☐ Yes ☐ No |  |
|  | Pressure regulated volume control (PRVC) or similar . |  | ☐ Yes ☐ No |  |
|  | Non-invasive ventilation (CPAP / BiPAP). |  | ☐ Yes ☐ No |  |
|  | Apnea-backup ventilation. |  | ☐ Yes ☐ No |  |
|  | HIFLOW O2 |  | ☐ Yes ☐ No |  |
|  | WEANING MODE |  | ☐ Yes ☐ No |  |
|  | Monitored and controlled parameters |  | ☐ Yes ☐ No |  |
|  | FiO2: 21–100%. |  | ☐ Yes ☐ No |  |
|  | Tidal Volume: 5 - 2,000 mL |  | ☐ Yes ☐ No |  |
|  | Inspiratory flow: 0 - 180 [L/min]. |  | ☐ Yes ☐ No |  |
|  | Inspiratory pressure: 0 – 40 [cmH2O] |  | ☐ Yes ☐ No |  |
|  | I:E ratio: 1:9 to 4:1 |  | ☐ Yes ☐ No |  |
|  | Respiratory Rate: 10–100 breaths/min |  | ☐ Yes ☐ No |  |
|  | Inspiratory pause |  | ☐ Yes ☐ No |  |
|  | Internal PEEP capability/range: 0–40 cmH2O |  | ☐ Yes ☐ No |  |
|  | Displayed parameters |  | ☐ Yes ☐ No |  |
|  | Data Displayed |  | ☐ Yes ☐ No |  |
|  | Trends |  | ☐ Yes ☐ No |  |
|  | Status indicators for ventilator mode, battery status, patient data, alarm settings. |  | ☐ Yes ☐ No |  |
|  | FiO2. |  | ☐ Yes ☐ No |  |
|  | Co2 |  | ☐ Yes ☐ No |  |
|  | Airway pressures (peak, plateau mean and PEEP). |  | ☐ Yes ☐ No |  |
|  | Tidal volume (inspired and expired). |  | ☐ Yes ☐ No |  |
|  | Minute volume (inspired and expired). |  | ☐ Yes ☐ No |  |
|  | I:E ratio |  | ☐ Yes ☐ No |  |
|  | RR (spontaneous and mechanical). |  | ☐ Yes ☐ No |  |
|  | Alarms, related to gas delivered |  | ☐ Yes ☐ No |  |
|  | Adjustable, visual and audible for: |  | ☐ Yes ☐ No |  |
|  | High/low FiO2. |  | ☐ Yes ☐ No |  |
|  | High/low inspiratory pressure and PEEP. |  | ☐ Yes ☐ No |  |
|  | High/low tidal volume (not achieved or exceeded). |  | ☐ Yes ☐ No |  |
|  | Apnea. |  | ☐ Yes ☐ No |  |
|  | High/low RR. |  | ☐ Yes ☐ No |  |
|  | Continuously high pressure/occlusion. |  | ☐ Yes ☐ No |  |
|  | Breathing circuit disconnect. |  | ☐ Yes ☐ No |  |
|  | Low minute volume. |  | ☐ Yes ☐ No |  |
|  | Alarms, related to equipment operation |  | ☐ Yes ☐ No |  |
|  | Visual and audible for: |  | ☐ Yes ☐ No |  |
|  | Gas supply failure. |  | ☐ Yes ☐ No |  |
|  | Power failure. |  | ☐ Yes ☐ No |  |
|  | Low battery. |  | ☐ Yes ☐ No |  |
|  | Self-diagnostics failure alarm. |  | ☐ Yes ☐ No |  |
|  | Mounting |  | ☐ Yes ☐ No |  |
|  | Ventilator must be mounted on original mobile cart |  | ☐ Yes ☐ No |  |
|  | Cart must be with at least 4 swivelling castors, 2 with brakes |  | ☐ Yes ☐ No |  |
|  | Cart must be capable to hold ventilator with humidifier and all accessories |  | ☐ Yes ☐ No |  |
|  | Movable arm holder for supporting patient breathing circuit |  | ☐ Yes ☐ No |  |
|  | Supplied with |  | ☐ Yes ☐ No |  |
|  | Oxygen hose, 3 m, with DIN probe for oxygen wall outlet |  | ☐ Yes ☐ No |  |
|  | Air hose, 3 m, with DIN probe for medical air wall outlet |  | ☐ Yes ☐ No |  |
|  | 10x disposable adult heated breathing circuit (incl.water trap & bacterial filter) or water preamble circuit, with chamber |  | ☐ Yes ☐ No |  |
|  | 10x disposable pediatric heated breathing circuit (incl. water trap & bacterial filter) or water preamble circuit, with chamber |  | ☐ Yes ☐ No |  |
|  | 15x disposable non-invasive ventilation masks, for adult (different sizes). |  | ☐ Yes ☐ No |  |
|  | 5x disposable non-invasive ventilation masks, for paediatric |  | ☐ Yes ☐ No |  |
|  | Test lungs for adult & pediatric applications |  | ☐ Yes ☐ No |  |
|  | Servo-controlled Humidifier with temprature probe and heated wire, the humidifier should have two modes (invasive temp 37) & (non invasive temp 31), with audiable and visable alarm |  | ☐ Yes ☐ No |  |
|  | ALL STANDARD ACCESSORIES TO MAKE UNIT FULLY FUNCTINAL & IMMEDIATE USE |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | Compliant with IEC 60601 or equivalent |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |
| 2 | Bedhead unit Horizontal, ICU |  |  |  |
|  | Name of Manufacturer | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | The supplier shall supply and install a horizontally wall-mounted patient bedhead service unit / trunking, complete with provisions for power supply, nurse call, medical gases and other facilities. |  | ☐ Yes ☐ No |  |
|  | "It shall be constructed of extruded anodized aluminum alloy sections. Anodized aluminum facials shall be removable to provide access to individual components mounted within the unit. Front panels shall be finished with high-pressure plastic laminate face." |  | ☐ Yes ☐ No |  |
|  | "The Horizontal Bedhead Unit shall comprise the services outlined on the corresponding detailed loaded drawings (1:50) and consisting of: 1. Five x duplix electric socket outlet switched 2. Two Double Data outlet 3. Nurse Call system 4. Medical gases outlets: Two medical oxygen outlet,Two medical vacuum outlet, One medical compressed air (4 bar) outlet direct Light & indirect light with On/Off switch on the bed head unit" |  | ☐ Yes ☐ No |  |
|  | The outlets shall be Color coded. Coding to be coordinated with the Hospital’s electromechanical consultant or contractor |  | ☐ Yes ☐ No |  |
|  | The outlets shall be ergonomically spaced so that they can comfortably carry items such as oxygen regulators with humidifiers and suction regulators with suction jars without colliding or obstructing other service outlets. Connection of all outlets and connectors to the supply systems are the responsibility of the supplier and shall be coordinated wit the Hospital’s electro-mechanical consultant or contractor. |  | ☐ Yes ☐ No |  |
|  | The bedhead unit shall include accessory tracks for mounting various accessories. |  | ☐ Yes ☐ No |  |
|  | The bedhead unit installation and design shall meet the international quality and safety standards for medical gases, pipeline system and bedhead units adopted by the hospital. |  | ☐ Yes ☐ No |  |
|  | Approximate overall dimensions (D x H): 10 x 35 cm. For approximate Width dimension,PLEASE SPECIFY |  | ☐ Yes ☐ No |  |
|  | Shop drawings shall show dimensions and details of related services including wiring and piping diagrams. They shall also show a complete list of components, attachments and accessories for approval. |  | ☐ Yes ☐ No |  |
|  | The bedhead unit installation and design shall meet the standards and norms adopted by the electro-mechanical consultant (NFPA, HTM, etc.). |  | ☐ Yes ☐ No |  |
|  | Suppliers shall install medical gas outlet terminals and electrical socket outlets using the same type already installed in the existing hospital. |  | ☐ Yes ☐ No |  |
|  | Outlet provisions and outlet installations for data and nurse call systems shall be fully coordinated with the electromechanical consultant. |  | ☐ Yes ☐ No |  |
|  | The bedhead unit shall be pre-piped and pre-wired; factory assembled electrical and mechanical components. It shall also be factory tested and delivered to site with singular terminal connections. |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source. Compliant with IEC 60601 or equivalent. |  | ☐ Yes ☐ No |  |
| 3 | Bedhead unit Horizontal, Patient Room |  |  |  |
|  | Name of Manufacturer | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Each bed head unit consists of:2 general socket outlets |  | ☐ Yes ☐ No |  |
|  | 2 UPS socket outlet |  | ☐ Yes ☐ No |  |
|  | 1 Data socket outlet |  | ☐ Yes ☐ No |  |
|  | 1 Nurse call provision with reset push button. |  | ☐ Yes ☐ No |  |
|  | 1 Doctor call push button. |  | ☐ Yes ☐ No |  |
|  | 1 switch for inspection lamp |  | ☐ Yes ☐ No |  |
|  | 1 Earthing outlet |  | ☐ Yes ☐ No |  |
|  | 1 Oxygen outlet points DIN type. |  | ☐ Yes ☐ No |  |
|  | 1 Vacuum outlet point DIN type. |  | ☐ Yes ☐ No |  |
|  | The bed head units length , should be according to the wall length as shown on the drawings. |  | ☐ Yes ☐ No |  |
|  | Units installation shall be in full compliance with the HTM 02-01.ISO 9001, ISO 13485 or equivalent |  | ☐ Yes ☐ No |  |
| 4 | Bedhead unit Vertical, Patient Room |  |  |  |
|  | Name of Manufacturer | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Each bed head unit consists of: 2 general socket outlets |  | ☐ Yes ☐ No |  |
|  | 2 UPS socket outlet |  | ☐ Yes ☐ No |  |
|  | 1 Data socket outlet |  | ☐ Yes ☐ No |  |
|  | 1 Nurse call provision with reset push button |  | ☐ Yes ☐ No |  |
|  | 1 Doctor call push button |  | ☐ Yes ☐ No |  |
|  | 1 switch for inspection lamp |  | ☐ Yes ☐ No |  |
|  | 1 Earthing outlet |  | ☐ Yes ☐ No |  |
|  | 1 Oxygen outlet points DIN type |  | ☐ Yes ☐ No |  |
|  | 1 Vacuum outlet point DIN type. |  | ☐ Yes ☐ No |  |
|  | The bed head units length , should be according to the wall length as shown on the drawings. |  | ☐ Yes ☐ No |  |
|  | ISO 9001, ISO 13485 or equivalent" |  | ☐ Yes ☐ No |  |
| 5 | Defibrillator Biphasic Adult/Pediatric |  |  |  |
|  | Manufacturer | 2 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Defibrillator for adult and pediatric patients, with ECG, NIBP and SPo2 monitoring and a printer, synchronized, AED function and bi-phasic, AC and battery powered, with accessories |  | ☐ Yes ☐ No |  |
|  | Light and sturdy portable unit |  | ☐ Yes ☐ No |  |
|  | Operation modes: Manual, Synchronized and AED |  | ☐ Yes ☐ No |  |
|  | Monitoring function mode parameters, with standard alarms: |  | ☐ Yes ☐ No |  |
|  | ECG |  | ☐ Yes ☐ No |  |
|  | HR |  | ☐ Yes ☐ No |  |
|  | NIBP |  | ☐ Yes ☐ No |  |
|  | SPo2 |  | ☐ Yes ☐ No |  |
|  | Colored LCD 7 Inch display, with intuitive controls and selectors Easy 1-2-3 utilization set up |  | ☐ Yes ☐ No |  |
|  | Displayed information: HR, ECG waveform, source, SPO2, NIBP and energy selected and given |  | ☐ Yes ☐ No |  |
|  | Bi-phasic defibrillation wave |  | ☐ Yes ☐ No |  |
|  | Max. Energy output : ≥200 J |  | ☐ Yes ☐ No |  |
|  | Discharge energy adjustable settings with multi step selector, from defibrillator body . |  | ☐ Yes ☐ No |  |
|  | Settable ECG synchronized discharges for cardio-version |  | ☐ Yes ☐ No |  |
|  | Charging indication |  | ☐ Yes ☐ No |  |
|  | Facility for self-test/check before usage and set up function |  | ☐ Yes ☐ No |  |
|  | Built in print recorder with thermal head |  | ☐ Yes ☐ No |  |
|  | ECG Source: Manually selected, paddles, leads: I, II, III |  | ☐ Yes ☐ No |  |
|  | Supplied with accessories |  | ☐ Yes ☐ No |  |
|  | One (1) ECG reusable patient cable with 3 leads , defibrillation protected |  | ☐ Yes ☐ No |  |
|  | Three (3) set of 50 disposable ECG adhesive electrodes. |  | ☐ Yes ☐ No |  |
|  | 20 pairs disposable patches (adult and pediatric) for semiautomatic mode, with related cable connection to the defibrillator |  | ☐ Yes ☐ No |  |
|  | Set of five (5) rolls of thermal paper. |  | ☐ Yes ☐ No |  |
|  | One (1) set of reusable adult external paddles (apex-sternum) and related pediatric adaptors compatible with the equipment |  | ☐ Yes ☐ No |  |
|  | One (1) NIPB cuff adult with hose |  | ☐ Yes ☐ No |  |
|  | One (1) NIPB cuff pediatric with hose |  | ☐ Yes ☐ No |  |
|  | One (1) Reusable SPO2 pediatric finger probe |  | ☐ Yes ☐ No |  |
|  | One (1) Reusable SPO2 adult finger probe |  | ☐ Yes ☐ No |  |
|  | Internal rechargeable battery with Low battery visual alarm |  | ☐ Yes ☐ No |  |
|  | ALL STANDARD ACCESSORIES TO MAKE UNIT FULLY FUNCTINAL & IMMEDIATE USE |  | ☐ Yes ☐ No |  |
|  | Rechargeable batteries |  | ☐ Yes ☐ No |  |
|  | Automatic switch from electric-line mode to battery operating mode |  | ☐ Yes ☐ No |  |
|  | Continuous monitoring working time in battery operating mode not less than 2 hours. |  | ☐ Yes ☐ No |  |
|  | Integrated battery charger. |  | ☐ Yes ☐ No |  |
|  | Low battery visual alarm |  | ☐ Yes ☐ No |  |
|  | 100% high capacity batteries with re-charging time not greater than 5 hours. |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |
| 6 | Patient Monitoring System (Bedside Monitors + Central Station) |  |  |  |
|  | Manufacturer | 4 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Country where the manufacturer is based |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting from the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Minimum UNOPS requirements |  | ☐ Yes ☐ No |  |
|  | CENTRAL STATION |  | ☐ Yes ☐ No |  |
|  | One (1) Central Station per system (not per monitor) |  | ☐ Yes ☐ No |  |
|  | Display Monitor : High resolution color not less than 21 Inch , Flat screen TFT |  | ☐ Yes ☐ No |  |
|  | Number Of Monitors:Two screen with configurable Number of patient |  | ☐ Yes ☐ No |  |
|  | Trending Full Disclosure:Not Less Than 72 Hrs |  | ☐ Yes ☐ No |  |
|  | Communication:Network, central station, remote bedside monitor |  | ☐ Yes ☐ No |  |
|  | Configurable Reporting |  | ☐ Yes ☐ No |  |
|  | keyboard and mouse. |  | ☐ Yes ☐ No |  |
|  | Laser Printer |  | ☐ Yes ☐ No |  |
|  | All patients data |  | ☐ Yes ☐ No |  |
|  | Waveforms |  | ☐ Yes ☐ No |  |
|  | Alarm - Audio and Visible |  | ☐ Yes ☐ No |  |
|  | Alarm Lamp and flashing |  | ☐ Yes ☐ No |  |
|  | Central Station Must Be Compatible with the selected Bed Side Monitor |  | ☐ Yes ☐ No |  |
|  | Patient Monitor |  | ☐ Yes ☐ No |  |
|  | Modular patient monitor with ICU software |  | ☐ Yes ☐ No |  |
|  | ECG,NIBP, SpO2, RR, TEMP, w/ accessories |  | ☐ Yes ☐ No |  |
|  | Internal, easily accessible battery lasting for at least 2 hours |  | ☐ Yes ☐ No |  |
|  | Wall mounting accessories with basket (original) supplied with the monitor |  | ☐ Yes ☐ No |  |
|  | Display: |  | ☐ Yes ☐ No |  |
|  | a)High resolution multi-color display (LCD) or TFT |  | ☐ Yes ☐ No |  |
|  | b)Touch, Size not less than 15" |  | ☐ Yes ☐ No |  |
|  | Simultaneously displayed information on the same screen: Numeric measured data and waveforms |  | ☐ Yes ☐ No |  |
|  | Visual and audible alarms with adjustable volume and limits |  | ☐ Yes ☐ No |  |
|  | PARAMETERS TO BE MONITORED: |  | ☐ Yes ☐ No |  |
|  | 1-ECG 10 LEAD, 5 LEAD, 3 LEAD with alarms |  | ☐ Yes ☐ No |  |
|  | 2-Respiration rate with alarms (Hi & low R.R. and apnea alarm) |  | ☐ Yes ☐ No |  |
|  | 3-NIBP: Systolic, Diastolic & mean – Mode: Automatic & Manual with alarms |  | ☐ Yes ☐ No |  |
|  | 4-Body temperature with alarms |  | ☐ Yes ☐ No |  |
|  | 5-SpO2 measurement with alarms |  | ☐ Yes ☐ No |  |
|  | 6-Heart Rate with alarms |  | ☐ Yes ☐ No |  |
|  | Options: |  | ☐ Yes ☐ No |  |
|  | Massimo technology option with accessories |  | ☐ Yes ☐ No |  |
|  | 2-channels Ibp option with accessories |  | ☐ Yes ☐ No |  |
|  | CO2 sidestream / main stream option with accessories |  | ☐ Yes ☐ No |  |
|  | Cardiac output PICO invasive option with accessories |  | ☐ Yes ☐ No |  |
|  | TRENDING Hrs.Not less than 48 hrs. for all monitored parameters |  | ☐ Yes ☐ No |  |
|  | Protections of all the functions against defibrillator discharges and electrosurgical units. |  | ☐ Yes ☐ No |  |
|  | Enclosure to have ingress protection level IPX1 or better. |  | ☐ Yes ☐ No |  |
|  | Accessories required for the above parameters |  | ☐ Yes ☐ No |  |
|  | Each unit must come complete with the following original accessories: |  | ☐ Yes ☐ No |  |
|  | ECG cable (3 leads), (with ECG trunk cable); Qty 2 |  | ☐ Yes ☐ No |  |
|  | ECG cable (5 leads), (with ECG trunk cable); Qty 2 |  | ☐ Yes ☐ No |  |
|  | ECG cable (10 leads), (with ECG trunk cable); Qty 2 |  | ☐ Yes ☐ No |  |
|  | Reusable Adult SpO2 sensor (with connector cable), 3 meters long; Qty 2 |  | ☐ Yes ☐ No |  |
|  | Reusable Pediatric SPO2 sensor (with connector cable), 3 meters long; Qty 2 |  | ☐ Yes ☐ No |  |
|  | Reusable Neonate SPO2 sensor (with connector cable), 3 meters long; Qty 2 |  | ☐ Yes ☐ No |  |
|  | Temperature probe reusable, external skin contact type; Qty 2 |  | ☐ Yes ☐ No |  |
|  | NIBP Neonate cuff with hose; Qty 2 |  | ☐ Yes ☐ No |  |
|  | NIBP Pediatric cuff with hose; Qty 2 |  | ☐ Yes ☐ No |  |
|  | NIBP Adult cuff with hose; Qty 2 |  | ☐ Yes ☐ No |  |
|  | Connectivity: |  | ☐ Yes ☐ No |  |
|  | Network connectivity: Monitor is able to connect to a Central station |  | ☐ Yes ☐ No |  |
|  | HL7 compatibility |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source. Compliant with IEC 60601 or equivalent. |  | ☐ Yes ☐ No |  |
| 7 | ECG Machine 12 Channels with Trolley |  |  |  |
|  | Name of Manufacturer | 4 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Electrocardiogram (ECG) digital monitor and recorder, 12-leads detection, multi-channel recording, mobile, AC and battery powered, with printer and accessories |  | ☐ Yes ☐ No |  |
|  | 12 channels ECG |  | ☐ Yes ☐ No |  |
|  | Measurement, interpretation, arrhythmia detection and ECG analysis |  | ☐ Yes ☐ No |  |
|  | LCD Display for real time waves preview, electrodes placement, numeric data: ≥ 4 inches |  | ☐ Yes ☐ No |  |
|  | Keyboard integrated for patient data inserting |  | ☐ Yes ☐ No |  |
|  | Automatic, manual and urgent mode recording |  | ☐ Yes ☐ No |  |
|  | Recording speed should be 25 mm/ sec and 50 mm/ sec. |  | ☐ Yes ☐ No |  |
|  | Print output of different waves format (3x4, 6x2, 12x1) |  | ☐ Yes ☐ No |  |
|  | Printer should preferably to work with standard thermal paper (should be available in Local Market) |  | ☐ Yes ☐ No |  |
|  | Integrated high resolution thermal printer for waves, patient data, date time, hospital, interpretation |  | ☐ Yes ☐ No |  |
|  | Basic memory storage of exams |  | ☐ Yes ☐ No |  |
|  | Standard settings for filtering, patient data type, units’ selection, speed, waves gain |  | ☐ Yes ☐ No |  |
|  | With standard alarms |  | ☐ Yes ☐ No |  |
|  | Complete with trolley with : Castors with brakes An integrated basket or drawer for accessories. ECG Cable holder |  | ☐ Yes ☐ No |  |
|  | Equipment compatible with patients with pacemakers and defibrillation protection. |  | ☐ Yes ☐ No |  |
|  | Supplied with accessories |  | ☐ Yes ☐ No |  |
|  | ECG patient cable, with preferably detachable 12 leads electrodes extension |  | ☐ Yes ☐ No |  |
|  | 2 sets of 4 limbs clamps |  | ☐ Yes ☐ No |  |
|  | 2 sets of 6 section-cup type electrodes |  | ☐ Yes ☐ No |  |
|  | 2 Sets of 6 clips for disposable electrodes connection |  | ☐ Yes ☐ No |  |
|  | 200 disposable adhesive electrodes |  | ☐ Yes ☐ No |  |
|  | 2 bottles of conductivity enhancing gel for adhesive electrodes |  | ☐ Yes ☐ No |  |
|  | 4 sets of thermal paper roll |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  | ☐ Yes ☐ No |  |
|  | Internal rechargeable battery |  | ☐ Yes ☐ No |  |
|  | Automatic switch from electric-line to battery operating mode and vice versa. |  | ☐ Yes ☐ No |  |
|  | Should have a digital filter for AC and EMG |  | ☐ Yes ☐ No |  |
|  | Continuous monitoring working time in battery operating mode for not less than 1 hour. |  | ☐ Yes ☐ No |  |
| 8 | Ultrasonic Nebulizer with Trolley |  |  |  |
|  | Manufacturer | 4 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Heavy duty construction, continuous operation |  | ☐ Yes ☐ No |  |
|  | Autoclavable chamaber and patient tubes. |  | ☐ Yes ☐ No |  |
|  | Light weight |  | ☐ Yes ☐ No |  |
|  | Digital display of nebulizing time |  | ☐ Yes ☐ No |  |
|  | Flexible arm support (for aerosol hose) |  | ☐ Yes ☐ No |  |
|  | Flat touch panel to power on/off the unit adjust nebulizing time and intensity. |  | ☐ Yes ☐ No |  |
|  | Transducer frequency: 1.7 MHZ approximately . |  | ☐ Yes ☐ No |  |
|  | Antibacterial filter. |  | ☐ Yes ☐ No |  |
|  | Nebulizing capacity adjustable: 0-4.5 ml/min |  | ☐ Yes ☐ No |  |
|  | Nebulizing chamber capacity: not less than 500 ml |  | ☐ Yes ☐ No |  |
|  | One heat -able aerosol hose. |  | ☐ Yes ☐ No |  |
|  | Complete with original trolley on 5 castors |  | ☐ Yes ☐ No |  |
|  | All accessories inculding adult and pediatric accessories must be included in the offer. |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source Compliant with IEC 60601 or equivalent |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions and defibrillator discharge. |  | ☐ Yes ☐ No |  |
|  | Internal rechargeable battery |  | ☐ Yes ☐ No |  |
| 9 | Electrical ICU Bed 4 Sections with Mattress |  |  |  |
|  | Name of Manufacturer | 10 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Designed for ICUs to take care of critical patients. |  | ☐ Yes ☐ No |  |
|  | Four sections platform |  | ☐ Yes ☐ No |  |
|  | Electric Motors function, with at least 5 Functions of Hospital Bed (High/Low, Knee Raise, Back Raise, Trendelenburg and Reverse Trendelenburg) |  | ☐ Yes ☐ No |  |
|  | Radiolucent back section (head of bed) |  | ☐ Yes ☐ No |  |
|  | Heavy duty non-metallic removable Headboard and Footboard.. |  | ☐ Yes ☐ No |  |
|  | With split side rails articulated with the bed for full lateral protection in all positions |  | ☐ Yes ☐ No |  |
|  | With electrical function lock outs, simultaneous and individual |  | ☐ Yes ☐ No |  |
|  | Backrest and leg section movements are accompanied by auto-regression to reduce pressure on the abdomen. |  | ☐ Yes ☐ No |  |
|  | With 4 corner roller bumpers |  | ☐ Yes ☐ No |  |
|  | At least two IV-pole receptacles to be located at two sides of the bed |  | ☐ Yes ☐ No |  |
|  | One compatible IV pole with double hooks |  | ☐ Yes ☐ No |  |
|  | Bilateral lever for CPR control |  | ☐ Yes ☐ No |  |
|  | Slow dampened release side rails |  | ☐ Yes ☐ No |  |
|  | Length: 220,Width: 90 |  | ☐ Yes ☐ No |  |
|  | Minimum Height: 45-50cm |  | ☐ Yes ☐ No |  |
|  | Maximum Height: 75-80cm |  | ☐ Yes ☐ No |  |
|  | Head section articulation: 65 degrees, or higher |  | ☐ Yes ☐ No |  |
|  | Knee section articulation: 20 degrees, or higher |  | ☐ Yes ☐ No |  |
|  | Trendelenburg and Reverse Trendelenburg: 12 degrees, or higher |  | ☐ Yes ☐ No |  |
|  | Zero clearance between mattress and headboard |  | ☐ Yes ☐ No |  |
|  | Cardiac position one command on remote control |  | ☐ Yes ☐ No |  |
|  | Four heavy-duty medical grade , nonmarking castors of not less than 12 cm diameter.at least one electro-conductive castor. |  | ☐ Yes ☐ No |  |
|  | Central brake system with steering function |  | ☐ Yes ☐ No |  |
|  | Safe working load: at least 250 kg |  | ☐ Yes ☐ No |  |
|  | With built-in Battery Back up |  | ☐ Yes ☐ No |  |
|  | Liquid spill proof of IPX 4 degree or higher |  | ☐ Yes ☐ No |  |
|  | Integrated patient exit alarm |  | ☐ Yes ☐ No |  |
|  | With embedded integrated caregiver control panels on both sides of the bed (upper left and upper right) on the outer side of the rails. |  | ☐ Yes ☐ No |  |
|  | Remote control |  | ☐ Yes ☐ No |  |
|  | under bed clearance at least 10 cm |  | ☐ Yes ☐ No |  |
|  | scale system,Seperately Qoute |  | ☐ Yes ☐ No |  |
|  | Bed Mattress |  | ☐ Yes ☐ No |  |
|  | Medical grade mattress |  | ☐ Yes ☐ No |  |
|  | Must cover the whole sleeping surface platform |  | ☐ Yes ☐ No |  |
|  | Thickness: at least ~15 cm |  | ☐ Yes ☐ No |  |
|  | Patient weight capacity: up to 200kg or more |  | ☐ Yes ☐ No |  |
|  | Anti-microbial and anti-fungus |  | ☐ Yes ☐ No |  |
|  | Complete with fluid resistant, non-latex, breathable and washable cover. |  | ☐ Yes ☐ No |  |
|  | visco-elastic foam core |  | ☐ Yes ☐ No |  |
|  | Antistatic, Antibacterial |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | "A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions |  | ☐ Yes ☐ No |  |
| 10 | Mobile Electric Suction Machine |  |  |  |
|  | Manufacturer | 4 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Suction unit, (high vacuum/high flow), 2 bottles, with cart, AC powered |  | ☐ Yes ☐ No |  |
|  | The pump shall be robust, maintenance-free, for hospital use |  | ☐ Yes ☐ No |  |
|  | The pumped liquid shall be sealed off from the pump |  | ☐ Yes ☐ No |  |
|  | Maximum vacuum: not less than -640 mmHg |  | ☐ Yes ☐ No |  |
|  | Come with suction controller and vacuum gauge / indicator |  | ☐ Yes ☐ No |  |
|  | Come with overflow control valves and safety mechanisms |  | ☐ Yes ☐ No |  |
|  | Maximum suction rate shall be at least 40 l/min |  | ☐ Yes ☐ No |  |
|  | The pump shall come fitted with 2 autoclavable polycarbonate suction bottles, around 2 liters each |  | ☐ Yes ☐ No |  |
|  | The bottles shall be incorporated with a safety suction cut-off mechanism when they become full |  | ☐ Yes ☐ No |  |
|  | It should be fitted with bacterial hydrophobic filter |  | ☐ Yes ☐ No |  |
|  | Accessories & Consumables |  | ☐ Yes ☐ No |  |
|  | Complete connection tubing set: 1 set |  | ☐ Yes ☐ No |  |
|  | Foot switch with cables |  | ☐ Yes ☐ No |  |
|  | Bacterial filter: 10 pcs |  | ☐ Yes ☐ No |  |
|  | Spare autoclavable polycarbonate suction bottle 2L: 1pc |  | ☐ Yes ☐ No |  |
|  | Trolley (with fitting for bottles) on 4 castors |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase compliant with IEC 60601 or equivalent |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage, over-current line conditions |  | ☐ Yes ☐ No |  |
| 11 | Bed, Patient, Hydraulic |  |  |  |
|  | Name of Manufacturer | 21 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Designed: Heavy Duty Designed, New Model. |  | ☐ Yes ☐ No |  |
|  | The Machine Should be of Latest Model and Must have Lattest Technology |  | ☐ Yes ☐ No |  |
|  | Quality: High Quality, Safety, Ease of Use |  | ☐ Yes ☐ No |  |
|  | Design & Quality: Bed, designed 4 Sections standard wards . |  | ☐ Yes ☐ No |  |
|  | Construction: Epoxy-Coated main frame constructed from heavy duty section steel tube |  | ☐ Yes ☐ No |  |
|  | Removable Chrome-Plated steel bed ends fitted with 10mm laminated plastic panels |  | ☐ Yes ☐ No |  |
|  | 4 Sections mattress Platform fitted with Epoxy-Coated steel weldmesh |  | ☐ Yes ☐ No |  |
|  | Backrest and knee brake adjustable by hydraulic action |  | ☐ Yes ☐ No |  |
|  | With split side rails articulated with the bed for full lateral protection in all positions |  | ☐ Yes ☐ No |  |
|  | With 4 corner roller bumpers |  | ☐ Yes ☐ No |  |
|  | At least two IV-pole receptacles to be located at two sides of the bed |  | ☐ Yes ☐ No |  |
|  | One compatible IV pole with double hooks |  | ☐ Yes ☐ No |  |
|  | Bilateral lever for CPR control |  | ☐ Yes ☐ No |  |
|  | Slow dampened release side rails |  | ☐ Yes ☐ No |  |
|  | Four heavy-duty medical grade electro-conductive, nonmarking castors of not less than 12 cm diameter. |  | ☐ Yes ☐ No |  |
|  | Central brake system with steering function |  | ☐ Yes ☐ No |  |
|  | Safe working load: at least 180 kg |  | ☐ Yes ☐ No |  |
|  | Guiding Dimentions of bed: 210-220L x 90-100Wx 50-80H cm |  | ☐ Yes ☐ No |  |
|  | Bed Mattress |  | ☐ Yes ☐ No |  |
|  | Medical grade mattress |  | ☐ Yes ☐ No |  |
|  | Must cover the whole sleeping surface platform |  | ☐ Yes ☐ No |  |
|  | Three sections foam. Mattress, 10-12cm at least High density, Premium-grade cotton felt and high-density urethane foam are intertufted to eliminate component shifting and to increase mattress life with Durable vinyl cover, anti-bacterial, anti-static, acid-resistant and waterproof for easy use and care |  | ☐ Yes ☐ No |  |
|  | Patient weight capacity: up to180kg or more |  | ☐ Yes ☐ No |  |
|  | Anti-microbial and anti-fungus |  | ☐ Yes ☐ No |  |
|  | Complete with fluid resistant, non-latex, breathable and washable cover. |  | ☐ Yes ☐ No |  |
|  | 1 X Pillow With Cases |  | ☐ Yes ☐ No |  |
|  | 1 X Hospital Sheet |  | ☐ Yes ☐ No |  |
|  | Thickness: at least ~15 cm |  | ☐ Yes ☐ No |  |
|  | Anti-microbial and anti-fungus |  | ☐ Yes ☐ No |  |
|  | Safety and standards |  | ☐ Yes ☐ No |  |
|  | ISO 13485 or 9001 certification of the Manufacturer’s QMS, issued by EU Notified Bodies or by an IAF recognized/accredited Conformity Assessment Body |  | ☐ Yes ☐ No |  |
|  | Copy of the above mentioned certificates shall be included in the offer |  | ☐ Yes ☐ No |  |
|  | The estimated life span : 8 -10 years |  | ☐ Yes ☐ No |  |
|  | The product to be supplied shall be new, unused and conform to the standards as specified in technical specifications. |  | ☐ Yes ☐ No |  |
|  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion. |  | ☐ Yes ☐ No |  |
|  | The equipment proposed shall be of highest quality and produced by well known manufacturers. The equipment shall carry the name and quality label of the manufacturer and fulfil the standards in force. |  | ☐ Yes ☐ No |  |
|  | Technical offers must include brochures, data sheets and technical complete technical specifications. |  | ☐ Yes ☐ No |  |
|  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number. |  | ☐ Yes ☐ No |  |
|  | All equipment to be brand new, and latest in design and technology. |  | ☐ Yes ☐ No |  |
|  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation. |  | ☐ Yes ☐ No |  |
|  | The staff training in operation and maintenance shall be provided |  | ☐ Yes ☐ No |  |
|  | Supplier should deliver with the machines all operator manuals, service manuals, engineering schematics, and all documents and software media relevant to the machines |  | ☐ Yes ☐ No |  |
|  | To submit a spare part pricelist with the part number valid for 7 years starting date after the warranty completion |  | ☐ Yes ☐ No |  |
| 12 | Pulse Oximeter (Handheld) |  |  |  |
|  | Manufacturer | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Application: High quality pulse oximeter that reads oxygen saturation to perform safe anesthesia and monitor patient in ICU & OR be |  | ☐ Yes ☐ No |  |
|  | Measurement range of O2 sat. (0-100) % |  | ☐ Yes ☐ No |  |
|  | Accuracy of O2 sat: ± 2% (70-100) |  | ☐ Yes ☐ No |  |
|  | Range of pulse rate: 30-240 bpm |  | ☐ Yes ☐ No |  |
|  | Display LCD or digital for all readings |  | ☐ Yes ☐ No |  |
|  | Visual and audible signal of pulse included |  | ☐ Yes ☐ No |  |
|  | Patient range: Adult, pediatric, infant, neonate. |  | ☐ Yes ☐ No |  |
|  | Alarms limits :(Adjustable) |  | ☐ Yes ☐ No |  |
|  | High/low pulse rate |  | ☐ Yes ☐ No |  |
|  | High and low spO2, Sensor off |  | ☐ Yes ☐ No |  |
|  | Battery low |  | ☐ Yes ☐ No |  |
|  | System failure |  | ☐ Yes ☐ No |  |
|  | All alarms audible and visible |  | ☐ Yes ☐ No |  |
|  | Alarm reset, silent |  | ☐ Yes ☐ No |  |
|  | Probe type (Quantity: 1 for each type) |  | ☐ Yes ☐ No |  |
|  | Reusable finger probe for Adult (length 3 m) |  | ☐ Yes ☐ No |  |
|  | Reusable finger probe for pediatric (length 3 m) |  | ☐ Yes ☐ No |  |
|  | Reusable probe for neonate (length 3 m) |  | ☐ Yes ☐ No |  |
|  | All needed accessories: Included |  | ☐ Yes ☐ No |  |
|  | Rechargeable battery with charger: Built in |  | ☐ Yes ☐ No |  |
|  | Power supply: 220 V / 50 Hz |  | ☐ Yes ☐ No |  |
| 13 | Medical Oxygen Concentrator |  |  |  |
|  | Manufacturer | 4 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Country where the manufacturer is based |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting from the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | The Oxygen Concentrator Shall Feature a One Head Compressor. |  | ☐ Yes ☐ No |  |
|  | Producing. Capacity: 1 to 10 L/Min |  | ☐ Yes ☐ No |  |
|  | Purity: 93+/- 3% , |  | ☐ Yes ☐ No |  |
|  | Outlet pressure (Mpa): At least 0.02-0.08 |  | ☐ Yes ☐ No |  |
|  | Pressure-compensated Flowmeter |  | ☐ Yes ☐ No |  |
|  | accurate flow setting. |  | ☐ Yes ☐ No |  |
|  | Safety Alarms: Audible and visual |  | ☐ Yes ☐ No |  |
|  | Power Failure. |  | ☐ Yes ☐ No |  |
|  | Restricted Flow. |  | ☐ Yes ☐ No |  |
|  | Low O2. |  | ☐ Yes ☐ No |  |
|  | Pressure-relief Valve and thermal protection of the Compressor |  | ☐ Yes ☐ No |  |
|  | Double-insulated Unit |  | ☐ Yes ☐ No |  |
|  | Flame-retardant Cabinet |  | ☐ Yes ☐ No |  |
|  | Sound Level (ANSI): 50 dB average |  | ☐ Yes ☐ No |  |
|  | Fixed humidifier Port and Recess shall prevent bottle and connector breakage |  | ☐ Yes ☐ No |  |
|  | Accessories: |  | ☐ Yes ☐ No |  |
|  | O2 Tubing. |  | ☐ Yes ☐ No |  |
|  | 3 X Face Masks ( Adult , Infant , New Born ). |  | ☐ Yes ☐ No |  |
|  | Humidifier. |  | ☐ Yes ☐ No |  |
|  | Set of Filters |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | Compliant with IEC 60601 or equivalent |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |
| 14 | Resuscitation Set (Adult & Child) |  |  |  |
|  | Manufacturer | 4 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Reusable hand-operated double-wall self-inflating device used to provide or assist pulmonary resuscitation. |  | ☐ Yes ☐ No |  |
|  | Clear, flexible cushion face mask allows for easy patient visualization |  | ☐ Yes ☐ No |  |
|  | Textured bag for a sure, comfortable grip |  | ☐ Yes ☐ No |  |
|  | Non-rebreathing valve that directs exhaled gas away from clinician |  | ☐ Yes ☐ No |  |
|  | Swivel patient connector |  | ☐ Yes ☐ No |  |
|  | Latex Free, made of Silicone |  | ☐ Yes ☐ No |  |
|  | Integrated PEEP connection, PEEP valve available separately |  | ☐ Yes ☐ No |  |
|  | Manual resuscitator with oxygen tubing, reservoir bag and pressure limiting valve. With mask, With PEEP valve |  | ☐ Yes ☐ No |  |
|  | Adult, Pediatric (Child), Infant Resuscitation bag |  | ☐ Yes ☐ No |  |
|  | Content: |  | ☐ Yes ☐ No |  |
|  | Resuscitator, PVC mask, oxygen reservoir bag, oxygen tubing 2.1M, pressure relief valve, PEEP valve. |  | ☐ Yes ☐ No |  |
| 15 | Electrosurgical Unit, Bipolar |  |  |  |
|  | Manufacturer | 5 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Original Electrosurgical Unit with cart |  | ☐ Yes ☐ No |  |
|  | A microcontroller based, isolated Electrosurgical Generator, having both monopolar and bipolar outputs designed for all surgical procedures |  | ☐ Yes ☐ No |  |
|  | Smart generator should be able to monitor changes in tissue impedance continuously and adjusts power |  | ☐ Yes ☐ No |  |
|  | Monopolar and bipolar functionalities |  | ☐ Yes ☐ No |  |
|  | Hand and pedal activation |  | ☐ Yes ☐ No |  |
|  | Power output: ≥ 300 Watts (Monopolar cut) |  | ☐ Yes ☐ No |  |
|  | The unit should have individual digital display of power for Bipolar, Monopolar cut and Monopolar Coag. |  | ☐ Yes ☐ No |  |
|  | Control panel allowing selection of modes and power setting |  | ☐ Yes ☐ No |  |
|  | Pre-existing range of cut and coagulation programs |  | ☐ Yes ☐ No |  |
|  | Audio-visual indicators of operation, alarms and errors |  | ☐ Yes ☐ No |  |
|  | The equipment shall provide at least the following standard modes (or equivalent): |  | ☐ Yes ☐ No |  |
|  | Monopolar Cut: Pure, Blend |  | ☐ Yes ☐ No |  |
|  | Monopolar Coag: Spray, soft/ desiccate, forced/ fulgurate |  | ☐ Yes ☐ No |  |
|  | Bipolar: Cut, Coag |  | ☐ Yes ☐ No |  |
|  | Safety System: |  | ☐ Yes ☐ No |  |
|  | It should have patient plate monitoring (Return Electrode Monitoring) |  | ☐ Yes ☐ No |  |
|  | Should give audiovisual alarm and deactivate output if contact between patient and patient plate is not proper |  | ☐ Yes ☐ No |  |
|  | Input sockets for monopolar & bipolar |  | ☐ Yes ☐ No |  |
|  | Mounted on a 4 wheels cart, antistatic, with brakes and space for accessories lodging (shelf / drawer) |  | ☐ Yes ☐ No |  |
|  | Supplied with accessories |  | ☐ Yes ☐ No |  |
|  | Monopolar Footswitch (Cut & Coag) |  | ☐ Yes ☐ No |  |
|  | Bipolar Footswitch |  | ☐ Yes ☐ No |  |
|  | 2x reusable neutral plate adult, complete with cable and standard connector |  | ☐ Yes ☐ No |  |
|  | 1x reusable neutral plate pediatric, complete with cable and standard connector |  | ☐ Yes ☐ No |  |
|  | 5x sterilizable electrode handles, with finger buttons cut-coag (finger switch controlled), with standard connector and connecting cable |  | ☐ Yes ☐ No |  |
|  | 5x sterilizable handle, without finger buttons, for pedal activation (foot switch controlled), with standard connector and connecting cable |  | ☐ Yes ☐ No |  |
|  | 5x set of 5 standard serializable blade electrodes of different standard sizes |  | ☐ Yes ☐ No |  |
|  | 5x set of standard special electrodes (ball, sphere, loop…) |  | ☐ Yes ☐ No |  |
|  | 2x bipolar pliers/ forceps, of different standard sizes complete with connection cable |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical. |  | ☐ Yes ☐ No |  |
|  | Compliant with IEC 60601 or equivalent |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |
| 16 | Monitor, Physiologic, Anesthesia |  |  |  |
|  | Manufacturer | 4 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Multi-parameter patient monitor for critical care patients for adult ,pediatric and Neonate use |  | ☐ Yes ☐ No |  |
|  | Continuous display on screen of patient ECG, respiration rate and heart rate, invasive and non-invasive blood pressure, body temperature, SpO2 and CO2 mainstream or side stream. |  | ☐ Yes ☐ No |  |
|  | Dynamic digital display that can show all active parameters. |  | ☐ Yes ☐ No |  |
|  | Unwanted parameters can be deselected from display. |  | ☐ Yes ☐ No |  |
|  | Operates from mains voltage (with no external transformers) or from internal rechargeable battery. |  | ☐ Yes ☐ No |  |
|  | Display: |  | ☐ Yes ☐ No |  |
|  | High resolution multi-color display (LCD) or TFT |  | ☐ Yes ☐ No |  |
|  | Size not less than 15" |  | ☐ Yes ☐ No |  |
|  | Simultaneously displayed information on the same screen: Numeric measured data and waveforms |  | ☐ Yes ☐ No |  |
|  | Visual and audible alarms with adjustable volume and limits |  | ☐ Yes ☐ No |  |
|  | User friendly |  | ☐ Yes ☐ No |  |
|  | Heavy duty with robust, easy to clean design |  | ☐ Yes ☐ No |  |
|  | Design: Modular |  | ☐ Yes ☐ No |  |
|  | Parameters: |  | ☐ Yes ☐ No |  |
|  | ECG (ECG Lead waveform, ECG lead label) with alarms(Leads off, Hi + Low HR, Atrium and ventricle tachycardia, Atrium, and ventricle fibrillation) |  | ☐ Yes ☐ No |  |
|  | Selection 3 or 6 or 12 leads ECG |  | ☐ Yes ☐ No |  |
|  | Values (ECG lead waveform, label, HR gain as min) |  | ☐ Yes ☐ No |  |
|  | Gain (5, 10, 20 mm/mV) |  | ☐ Yes ☐ No |  |
|  | ST, Arrhythmia, cascade ECG |  | ☐ Yes ☐ No |  |
|  | Respiration rate with alarms (Hi & low R.R. and apnea alarm) |  | ☐ Yes ☐ No |  |
|  | Respiration rate measurement range at least 0–100 bpm |  | ☐ Yes ☐ No |  |
|  | Minimum gradation 1 bpm. |  | ☐ Yes ☐ No |  |
|  | NIBP: Systolic, Diastolic & mean – Mode: Automatic & Manual with alarms |  | ☐ Yes ☐ No |  |
|  | Blood pressure monitoring range at least 30–270 mmHg |  | ☐ Yes ☐ No |  |
|  | Minimum gradation 1 mmHg. |  | ☐ Yes ☐ No |  |
|  | Internal pump for cuff inflation for non-invasive blood pressure measurement, with over pressure protection |  | ☐ Yes ☐ No |  |
|  | CO2 monitoring capabilities; mainstream or side stream. |  | ☐ Yes ☐ No |  |
|  | 2xIBP (IBP+ CVP) with alarms |  | ☐ Yes ☐ No |  |
|  | Body temperature with alarms |  | ☐ Yes ☐ No |  |
|  | Temperature range at least 30–40 °C |  | ☐ Yes ☐ No |  |
|  | Minimum gradation 0.1 °C |  | ☐ Yes ☐ No |  |
|  | Temperature probe to be reusable, external skin contact type |  | ☐ Yes ☐ No |  |
|  | SpO2 measurement with alarms- Massimo technology to be quoted as optional |  | ☐ Yes ☐ No |  |
|  | SpO2 measurement range at least 70–99% |  | ☐ Yes ☐ No |  |
|  | With accuracy better than ± 3% |  | ☐ Yes ☐ No |  |
|  | Minimum gradation 1%. |  | ☐ Yes ☐ No |  |
|  | Heart Rate with alarms |  | ☐ Yes ☐ No |  |
|  | Heart rate measurement range to be at least 30–250 bpm |  | ☐ Yes ☐ No |  |
|  | With accuracy better than ± 5 bpm |  | ☐ Yes ☐ No |  |
|  | Minimum gradation 1 bpm. |  | ☐ Yes ☐ No |  |
|  | Protections of all the functions against defibrillator discharges and electrosurgical units. |  | ☐ Yes ☐ No |  |
|  | Enclosure to have ingress protection level IPX1 or better. |  | ☐ Yes ☐ No |  |
|  | Accessories |  | ☐ Yes ☐ No |  |
|  | Wall mounting accessories with basket (original) supplied with the monitor |  | ☐ Yes ☐ No |  |
|  | Internal, easily accessible battery lasting for at least 2 hours |  | ☐ Yes ☐ No |  |
|  | Each unit must come complete with the following original accessories: |  | ☐ Yes ☐ No |  |
|  | ECG cable (3 leads), (with ECG trunk cable); Qty 2 |  | ☐ Yes ☐ No |  |
|  | ECG cable (5 leads), (with ECG trunk cable); Qty 2 |  | ☐ Yes ☐ No |  |
|  | Reusable Adult SpO2 sensor (with connector cable), 3 meters long; Qty 2 |  | ☐ Yes ☐ No |  |
|  | Reusable Pediatric SPO2 sensor (with connector cable), 3 meters long; Qty 2 |  | ☐ Yes ☐ No |  |
|  | Reusable Neonate SPO2 sensor (with connector cable), 3 meters long; Qty 2 |  | ☐ Yes ☐ No |  |
|  | If CO2 mainstream technology: Reusable tube adapter: Qty 2 & sensor: Qty If CO2 side stream technology: sample lines: Qty 20 & water traps: Qty 2 |  | ☐ Yes ☐ No |  |
|  | IBP kits with complete connection cables; Qty 1 |  | ☐ Yes ☐ No |  |
|  | Temperature probe; Qty 2 |  | ☐ Yes ☐ No |  |
|  | NIBP Neonate cuff with hose; Qty 2 |  | ☐ Yes ☐ No |  |
|  | NIBP Pediatric cuff with hose; Qty 2 |  | ☐ Yes ☐ No |  |
|  | NIBP Adult cuff with hose; Qty 2 |  | ☐ Yes ☐ No |  |
|  | Connectivity: Network connectivity: Monitor is able to connect to a Central station |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source. Compliant with IEC 60601 or equivalent. |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |
| 17 | Table, Surgical, Major |  |  |  |
|  | Manufacturer | 4 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Electromechanical or electro hydraulic mobile operating table column. |  | ☐ Yes ☐ No |  |
|  | Radiolucent Tabletop: capability to be used with C-Arm, no metal bars in the middle of the table top for clear image |  | ☐ Yes ☐ No |  |
|  | It shall have integrated, maintenance-free, rechargeable batteries with integrated battery charger. |  | ☐ Yes ☐ No |  |
|  | 4-section operating tabletop for general surgeries including the following: - Head section - back section - Seat section - Leg section |  | ☐ Yes ☐ No |  |
|  | Full function, corded hand control (wireless remote control to be offered as an option), the following functions must be included: - Height control (up & down) - Trendelenburg and reverse Trendelenburg - Tilt and back tilt - Flex & reflex - Leg Plates up and down - Zero position (return to level) - Table Lock and Unlock |  | ☐ Yes ☐ No |  |
|  | Override touch control panel with all remote control basic functions to be included and located on the table column for easy access. |  | ☐ Yes ☐ No |  |
|  | Maximum lifting weight capacity: 400 Kg or better. |  | ☐ Yes ☐ No |  |
|  | Approximate total length of the operating tabletop with all of its sections: Aprox 2000 mm approximately. |  | ☐ Yes ☐ No |  |
|  | Approximate tabletop width (without side rails): ~500 - 600 mm |  | ☐ Yes ☐ No |  |
|  | Height (without mattress):≤ 680 to ≥ 1050 mm. |  | ☐ Yes ☐ No |  |
|  | Trendelenburg / reverse Trendelenburg: 30° or better |  | ☐ Yes ☐ No |  |
|  | Lateral tilt: left / right: 20 ° or better |  | ☐ Yes ☐ No |  |
|  | Back plate: up / down: +80° / -40°. |  | ☐ Yes ☐ No |  |
|  | Leg plate downward:~ 90°. |  | ☐ Yes ☐ No |  |
|  | Fully motorized longitudinal shift of at least 400 mm and should be controlled from the remote control |  | ☐ Yes ☐ No |  |
|  | Motorized Leg plates controlled by the remote control in reverse position |  | ☐ Yes ☐ No |  |
|  | Electrical table lock and unlock triggered by the remote control. |  | ☐ Yes ☐ No |  |
|  | Table orientation: to work on normal and reverse orientations |  | ☐ Yes ☐ No |  |
|  | Heavy duty stainless steel construction of side-rails, column casing. |  | ☐ Yes ☐ No |  |
|  | Table frame should be made from heavy duty cast aluminium or stainless steel. |  | ☐ Yes ☐ No |  |
|  | It shall have stable base design with shock resistant base cover |  | ☐ Yes ☐ No |  |
|  | Four double antistatic swivel castors of at least 10 cm diameter with central braking system. |  | ☐ Yes ☐ No |  |
|  | It shall have the possibility to set up a multitude of patient positions for general procedures. |  | ☐ Yes ☐ No |  |
|  | Mattresses (complete set: head, upper back, lower back, pelvis, leg sections, arm boards) to be with at least the following specifications: - Visco-elastic memory Foam material: moulds to body shape - Latex free - Thickness at least 80 mm - Radio-translucent - Electrical conductive - Fire retardant - Anti-microbial - Solution resistant - Washable (easy to clean) - Seamless (cover should be full-welded) - waterproof - Black cover |  | ☐ Yes ☐ No |  |
|  | Side rails for attachment of accessories. |  | ☐ Yes ☐ No |  |
|  | All general surgery options and accessories shall be included in the offer and priced separately. |  | ☐ Yes ☐ No |  |
|  | For each table the following accessories and parts must be included(quoted separately): |  | ☐ Yes ☐ No |  |
|  | - Radial clamp joint (complete in stainless steel) Qty.6 Anaesthesia screen with clamp Qty.1 Infusion stand with clamp Qty.1 Instrument table with clamp Qty.1 Pair of Arm boards height adjustable Qty.1 Pair of Knee crutches (Goepel design) Qty.1 Pair of Lateral support Qty.1 Pair of Back Support Qty.1 i Pair of shoulder supports Qty.1 Body strap Qty.1 Accessories stand trolley Qty.1 |  | ☐ Yes ☐ No |  |
|  | Power Requirements: 240V +/- 10%, 50Hz, 1phase |  | ☐ Yes ☐ No |  |
| 18 | Monitor, Physiologic, Portable |  |  |  |
|  | Manufacturer | 4 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | The system must be offered by a manufacturing company and not by a supplier company. |  | ☐ Yes ☐ No |  |
|  | Portable type monitor |  | ☐ Yes ☐ No |  |
|  | External AC/DC power supply not accepted |  | ☐ Yes ☐ No |  |
|  | Adult, paediatric and neonatal patient monitor. |  | ☐ Yes ☐ No |  |
|  | Heavy duty with robust design. |  | ☐ Yes ☐ No |  |
|  | The following original accessories are required with each unit: |  | ☐ Yes ☐ No |  |
|  | One (3-leads) ECG cable. |  | ☐ Yes ☐ No |  |
|  | One (5-leads) ECG cable. |  | ☐ Yes ☐ No |  |
|  | Two Reusable adult SpO2 sensors. |  | ☐ Yes ☐ No |  |
|  | Two Reusable paediatric SpO2 sensors. |  | ☐ Yes ☐ No |  |
|  | Two Reusable neonatal SpO2 sensors. |  | ☐ Yes ☐ No |  |
|  | Two (adult) cuffs complete with hose. |  | ☐ Yes ☐ No |  |
|  | Two (paediatric) cuffs complete with hose. |  | ☐ Yes ☐ No |  |
|  | Two (neonate) cuffs complete with hose |  | ☐ Yes ☐ No |  |
|  | Two temperature sensors |  | ☐ Yes ☐ No |  |
|  | Ability to measure and display the following parameters: |  | ☐ Yes ☐ No |  |
|  | ECG wave form. |  | ☐ Yes ☐ No |  |
|  | NIBP. |  | ☐ Yes ☐ No |  |
|  | Body temperature. |  | ☐ Yes ☐ No |  |
|  | Heart rate. |  | ☐ Yes ☐ No |  |
|  | Weight: less than 4 kgm |  | ☐ Yes ☐ No |  |
|  | All mounting options to be available and quoted separately including wall mount, bench top and mobile trolley. |  | ☐ Yes ☐ No |  |
|  | Standard accessory sets must be itemized. |  | ☐ Yes ☐ No |  |
|  | Adjustable ECG gain. |  | ☐ Yes ☐ No |  |
|  | Baseline adjustment capability. |  | ☐ Yes ☐ No |  |
|  | High resolution display with sweeping motion. |  | ☐ Yes ☐ No |  |
|  | All data shall be in real time (waveform and numeric). |  | ☐ Yes ☐ No |  |
|  | The latest high end 12.1-inch colour TFT LCD screen |  | ☐ Yes ☐ No |  |
|  | Built-in rechargeable battery minimum life 3 hours of continuous monitoring |  | ☐ Yes ☐ No |  |
|  | Min 92 hours trend capacity |  | ☐ Yes ☐ No |  |
|  | Minimum of 4 waves display |  | ☐ Yes ☐ No |  |
|  | All available options and accessories must be offered and priced separately |  | ☐ Yes ☐ No |  |
|  | ne original compatible mobile stand on castors to be included in the offer. |  | ☐ Yes ☐ No |  |
|  | The capability to be connected to a central station |  | ☐ Yes ☐ No |  |
| 19 | Bassinet, LDRP |  |  |  |
|  | Name of Manufacturer | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | seamless clear ABS material bassinet with handgrips and name plate holder |  | ☐ Yes ☐ No |  |
|  | Flat and Trendelenburg positioning |  | ☐ Yes ☐ No |  |
|  | Approximate dimensions (L x W x H): 800x400x900mm |  | ☐ Yes ☐ No |  |
|  | Mattress with disinfectant tolerant, stain and flame resistant cover |  | ☐ Yes ☐ No |  |
|  | Polished tubular steel cart, scratch, rust and corrosion resistant |  | ☐ Yes ☐ No |  |
|  | Lockable top drawer |  | ☐ Yes ☐ No |  |
|  | Four swivel casters, two with breaks |  | ☐ Yes ☐ No |  |
| 20 | Bed, Birthing |  |  |  |
|  | Name of Manufacturer | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Delivery Bed Should have following essential specifications: |  | ☐ Yes ☐ No |  |
|  | Overall size 2000 – 2100 mm L , 900mm– 1000 mm W (with and without side rails up), 550 mm to 900 mm H (without mattress). |  | ☐ Yes ☐ No |  |
|  | Bed should have three section high pressure laminated top. |  | ☐ Yes ☐ No |  |
|  | It should have detachable laminated top for middle section and leg section for ease of cleaning. |  | ☐ Yes ☐ No |  |
|  | Leg section should be telescoped under backrest for lithotomy position. |  | ☐ Yes ☐ No |  |
|  | It should have control devise for making height and back adjustments.[ Manual as well as remote control]. |  | ☐ Yes ☐ No |  |
|  | Frame should be made of mild steel which should be pre treated and powder coated. |  | ☐ Yes ☐ No |  |
|  | Bed shall have swing down type side railings; railings shall be made from non-rusting polymer moulded material. |  | ☐ Yes ☐ No |  |
|  | It should have three sectional mattress and seat section should have large perineal cut. |  | ☐ Yes ☐ No |  |
|  | Mattress should be made of high resilience high quality P.U foam with stain resistant cover and zip. |  | ☐ Yes ☐ No |  |
|  | Should have a pair of bearing down patient hand grips with PVC covers and with location adjustability for patient convenience. |  | ☐ Yes ☐ No |  |
|  | All positions should be operated by electromechanical adjustment through a hand set; battery backup with inbuilt battery charger should be provided. |  | ☐ Yes ☐ No |  |
|  | The hand set should have indications for power on and the battery charge. |  | ☐ Yes ☐ No |  |
|  | It should have polymer molded easily detachable head and foot side panels. It should have four corner rubber buffers. |  | ☐ Yes ☐ No |  |
|  | Bed should be mounted on 12 cm – 15 cm non rusting castor wheels with locking mechanism. |  | ☐ Yes ☐ No |  |
|  | Castor should be made from high grade non floor staining synthetic material. |  | ☐ Yes ☐ No |  |
|  | Wheel center should have precision ball bearing to run smoothly. |  | ☐ Yes ☐ No |  |
|  | Should have infusion rods which have adjustable heights, quick release and attaches to all corners of bed. |  | ☐ Yes ☐ No |  |
|  | Should have a pair of upholstery aluminum crutches mounted on SS tubes. |  | ☐ Yes ☐ No |  |
|  | Should have sliding waste collection bowl at perineal part of table. |  | ☐ Yes ☐ No |  |
|  | It should be able to give Trendelenburg, reverse trendelburg and 60 - 70 degree sitting position both mechanically and electronically. |  | ☐ Yes ☐ No |  |
|  | The unit shall be CE marked and/or FDA approved |  | ☐ Yes ☐ No |  |
| 21 | Light, Exam/Procedure, Single, Ceiling |  |  |  |
|  | Name of Manufacturer | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Single head, ceiling mounted exam/procedure light |  | ☐ Yes ☐ No |  |
|  | LED technology to generate a cool and shadow-less light |  | ☐ Yes ☐ No |  |
|  | Maximum Luminance at 1m should be between 40,000 and 60,000 lux |  | ☐ Yes ☐ No |  |
|  | Light Intensity: adjustable from 50% to 100% of maximum output |  | ☐ Yes ☐ No |  |
|  | Color rendering index: not less than 93 |  | ☐ Yes ☐ No |  |
|  | Color temperature: 4,200 K approximately |  | ☐ Yes ☐ No |  |
|  | Field size diameter: 15 cm approximately |  | ☐ Yes ☐ No |  |
|  | Rotation: 180 ˚ |  | ☐ Yes ☐ No |  |
|  | The unit shall include an integrated control unit. |  | ☐ Yes ☐ No |  |
|  | Light head housing shall be sealed, dust and fracture proof |  | ☐ Yes ☐ No |  |
|  | The light head shall have a positioning handle |  | ☐ Yes ☐ No |  |
|  | All surfaces shall be resistant to disinfectants. |  | ☐ Yes ☐ No |  |
|  | The supplier shall supply and install the safest, most durable wall interface configuration and shall be coordinated and approved by the client and relevant contractors. |  | ☐ Yes ☐ No |  |
|  | The supplier shall coordinate with the client and electromechanical contractors for installation requirements, exact locations, wiring, power requirements…etc. |  | ☐ Yes ☐ No |  |
|  | Shall be CE marked and/or FDA approved |  | ☐ Yes ☐ No |  |
| 22 | Light, Phototherapy, Infant, Mobile |  |  |  |
|  | Name of Manufacturer | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Bulbs (halogen–quartz fluorescent tubes or light emitting diodes ) to provide, after filtering, light of wavelength approximately 420 to 480 nm. |  | ☐ Yes ☐ No |  |
|  | Adjustable light intensity. |  | ☐ Yes ☐ No |  |
|  | Hour meter showing total exposure time for current patient to be clearly visible by operator. |  | ☐ Yes ☐ No |  |
|  | Bilirubin total irradiance maximum (Ebi max): 2500μW/cm² |  | ☐ Yes ☐ No |  |
|  | Radiance uniformity: 0.4 or better |  | ☐ Yes ☐ No |  |
|  | Over temperature safety cut out to be included. |  | ☐ Yes ☐ No |  |
|  | Counter for lamp working hours and built-in timer for dose monitoring. |  | ☐ Yes ☐ No |  |
|  | Lamp replacement interval not less than 40000 hours or better. |  | ☐ Yes ☐ No |  |
|  | Light emission peak spectrum inside the range 400 - 500 nm. |  | ☐ Yes ☐ No |  |
|  | Mobile unit with 4 castor, at least 2 anti-static wheels and at least two brakes. |  | ☐ Yes ☐ No |  |
|  | Height Adjustable 1350mm - 1650mm |  | ☐ Yes ☐ No |  |
| 23 | Monitor, O.B., Intrapartum, Maternal/Fetal |  |  |  |
|  | Name of Manufacturer | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Fetal and maternal monitor for Antepartum/Intrapartum monitoring in labor & delivery rooms |  | ☐ Yes ☐ No |  |
|  | The unit shall include all required software, modules, transducers and accessories to monitor and record the following parameters (minimum): |  | ☐ Yes ☐ No |  |
|  | External Fetal Heart Rate (FHR): |  | ☐ Yes ☐ No |  |
|  | With two transducers, for simultaneous monitoring of twins |  | ☐ Yes ☐ No |  |
|  | The transducers shall be waterproof |  | ☐ Yes ☐ No |  |
|  | With coincidence detection between fetal heart rate and maternal heart rate |  | ☐ Yes ☐ No |  |
|  | Measuring range: 50 – 210 bpm |  | ☐ Yes ☐ No |  |
|  | Internal/Direct fetal ECG (DECG): |  | ☐ Yes ☐ No |  |
|  | Using spiral scalp electrode |  | ☐ Yes ☐ No |  |
|  | Measuring range: 50 – 210 bpm |  | ☐ Yes ☐ No |  |
|  | External Uterine Activity (UA): |  | ☐ Yes ☐ No |  |
|  | Toco transducer |  | ☐ Yes ☐ No |  |
|  | The transducer shall be waterproof |  | ☐ Yes ☐ No |  |
|  | With automatic re-zero |  | ☐ Yes ☐ No |  |
|  | Intrauterine pressure (IUP): |  | ☐ Yes ☐ No |  |
|  | With reusable pressure transducers |  | ☐ Yes ☐ No |  |
|  | With disposable catheters |  | ☐ Yes ☐ No |  |
|  | Maternal monitoring: |  | ☐ Yes ☐ No |  |
|  | ECG (MECG) |  | ☐ Yes ☐ No |  |
|  | SpO2 and Pulse Rate |  | ☐ Yes ☐ No |  |
|  | Non-Invasive Blood Pressure (NIBP) |  | ☐ Yes ☐ No |  |
|  | With color LCD display, size: 6” (minimum) |  | ☐ Yes ☐ No |  |
|  | With integrated multi-channel strip-chart recorder: |  | ☐ Yes ☐ No |  |
|  | Annotation of date, time, bed number, alarm status, and vital signs |  | ☐ Yes ☐ No |  |
|  | Automatic trace calibration and centering |  | ☐ Yes ☐ No |  |
|  | With networking capability to the central fetal monitoring and data management systems |  | ☐ Yes ☐ No |  |
|  | With networking capability to Hospital Information Systems (HIS) and Electronic Medical Records (EMR) via HL7 protocol |  | ☐ Yes ☐ No |  |
|  | With push-button event marker that allows the nurse/physician to note the timing of events |  | ☐ Yes ☐ No |  |
|  | It shall have internal backup memory allowing recovery of up to one hour of trace-data (preferable) |  | ☐ Yes ☐ No |  |
|  | With audio and visual alarming system |  | ☐ Yes ☐ No |  |
|  | The following accessories shall be included: |  | ☐ Yes ☐ No |  |
|  | Two Fetal Heart Rate ultrasound transducers |  | ☐ Yes ☐ No |  |
|  | Disposable fetal ECG scalp electrodes |  | ☐ Yes ☐ No |  |
|  | External Toco transducer, for external uterine activity |  | ☐ Yes ☐ No |  |
|  | Reusable intrauterine pressure transducer and connector cable |  | ☐ Yes ☐ No |  |
|  | Disposable Intrauterine pressure catheter |  | ☐ Yes ☐ No |  |
|  | Reusable belts for all transducers, and gel |  | ☐ Yes ☐ No |  |
|  | Recording paper pack |  | ☐ Yes ☐ No |  |
|  | Maternal ECG cable with leads |  | ☐ Yes ☐ No |  |
|  | Disposable maternal ECG electrodes |  | ☐ Yes ☐ No |  |
|  | Reusable adult SpO2 finger sensor |  | ☐ Yes ☐ No |  |
|  | Reusable adult NIBP cuff with hoses |  | ☐ Yes ☐ No |  |
|  | Shall be CE marked and/or FDA approved |  | ☐ Yes ☐ No |  |
| 24 | Patient Monitor, Physiologic (OR) |  |  |  |
|  | Name of Manufacturer | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | Should have TFT/ LCD/ Touchscreen display with at least 15 inches and at least 8 wave forms touch screen and numeric display simultaneously |  | ☐ Yes ☐ No |  |
|  | The waveform should be user selectable with preset user defined settings for varying care locations. |  | ☐ Yes ☐ No |  |
|  | Should be portable with carrying handle. |  | ☐ Yes ☐ No |  |
|  | Rechargeable Lithium-ion Battery |  | ☐ Yes ☐ No |  |
|  | Should have keys for quick access to main functions. |  | ☐ Yes ☐ No |  |
|  | Should be able to monitor ECG, SPO2, NIBP, 2 IBP, Respiration Rate, 2 temp & ETCO2, for adult, pediatric and neonatal patients as standard with adult, pediatric & Neonatal modes by default. IBP should incorporate SVV/ SPV/ PPV monitoring. |  | ☐ Yes ☐ No |  |
|  | 12 Lead ECG monitoring with full range of lethal and non-lethal arrhythmia recognition capability with detailed list in the monitor to set priority and ST analysis. |  | ☐ Yes ☐ No |  |
|  | Respiration & Apnea alarm. |  | ☐ Yes ☐ No |  |
|  | Manual, Auto and STAT mode for NIBP monitoring and ranges should be 30 to 230 mmHg. |  | ☐ Yes ☐ No |  |
|  | Pulse Oximeter (SPO2) with Plethysmograph & Pulse strength indicator With Variable pitch with change in SpO2. Should have low perfusion compensation and motion tolerance technology. Documentary proof shall be submitted. |  | ☐ Yes ☐ No |  |
|  | Mainstream / Side-stream / Microstream Capnography with display of CO2 wave form & digital values (ETCO2, FiCO2,RR. |  | ☐ Yes ☐ No |  |
|  | Should have separate volume control for beep sound for QRS and alarm sound. |  | ☐ Yes ☐ No |  |
|  | The display setting should have user defined setups variable as per applications for flexible use of the monitor in various clinical environments. |  | ☐ Yes ☐ No |  |
|  | Monitor should be able to connect with Central monitoring station & inter bed facility as standard feature with bidirectional communication. |  | ☐ Yes ☐ No |  |
|  | Should provide following accessories |  | ☐ Yes ☐ No |  |
|  | Mainstream / Micro stream / Side stream ETCO2 disposable kit for Adult - 25 nos, Paediatric/ Neonatal - 25 nos. |  | ☐ Yes ☐ No |  |
|  | 25 Nos of Disposable IBP transducers with all standard accessories & 6 nos of reusable adapter cable (type as requested by the end user). |  | ☐ Yes ☐ No |  |
|  | Reusable adult 5 lead ECG cable set (Clip type) – 2 nos. |  | ☐ Yes ☐ No |  |
|  | Reusable adult and pediatric SPO2 finger probes – 2 each, Extension cable -2 Nos. |  | ☐ Yes ☐ No |  |
|  | Disposable SPO2 probes for neonatal use - 2 Nos & Y sensors – 2 nos. |  | ☐ Yes ☐ No |  |
|  | NIBP cuffs for standard Adult- 5 Nos, Obese Adult - 2 Nos, Child and infant – 2 each. |  | ☐ Yes ☐ No |  |
|  | Temperature Probe -2Nos (esophageal / rectal/ surface). |  | ☐ Yes ☐ No |  |
|  | Should provide a facility/ net box to keep the accessories. |  | ☐ Yes ☐ No |  |
|  | Equipment performance should not be affected by electromagnetic radiation or conducted through power lines from another device. |  | ☐ Yes ☐ No |  |
|  | Should work on 200-240V AC/ 50Hz with inbuilt rechargeable battery. |  | ☐ Yes ☐ No |  |
|  | Should be compatible with HIS and should be HL7 compliant. |  | ☐ Yes ☐ No |  |
|  | CE marked and/or FDA approved |  | ☐ Yes ☐ No |  |
|  | Facility for last 24 hours full disclosure with wave forms trend review & facility for patient data entry. |  | ☐ Yes ☐ No |  |
|  | Monitor stand as per the following specification should be available |  | ☐ Yes ☐ No |  |
|  | Monitor stand - extruded Aluminum, powder coated. |  | ☐ Yes ☐ No |  |
|  | Should supply, install with necessary anchor fastners at the site. |  | ☐ Yes ☐ No |  |
|  | Monitor shall be fixed with stand securely. |  | ☐ Yes ☐ No |  |
|  | Should have facility to adjust height, tilt and turn |  | ☐ Yes ☐ No |  |
| 25 | Regulator, Suction, Intermittent/Continuous |  |  |  |
|  | Name of Manufacturer | 4 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | The regulator shall be break-resistant, durable and constructed of high quality materials |  | ☐ Yes ☐ No |  |
|  | The regulator shall have the following minimum features: |  | ☐ Yes ☐ No |  |
|  | Shall provide continuous operation modes |  | ☐ Yes ☐ No |  |
|  | Regulation range: 0 to -760 mmHg |  | ☐ Yes ☐ No |  |
|  | Flow rate: 0 - 80 l/min (approximately) |  | ☐ Yes ☐ No |  |
|  | Nominal working pressure: -760 mmHg |  | ☐ Yes ☐ No |  |
|  | With shutoff valve to prevent fluids from ingress to central pipeline system |  | ☐ Yes ☐ No |  |
|  | With bacteria filter |  | ☐ Yes ☐ No |  |
|  | With wall mounted reusable autoclavable collection jar (500 ml capacity) |  | ☐ Yes ☐ No |  |
|  | Directly connected to central vacuum system as per hospital standards. |  | ☐ Yes ☐ No |  |
|  | All surfaces shall be disinfectable |  | ☐ Yes ☐ No |  |
|  | The unit shall be CE marked and/or FDA approved |  | ☐ Yes ☐ No |  |
| 26 | Table, Surgical, Minor |  |  |  |
|  | Name of Manufacturer | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | The base attachment shall be a mobile type with castors |  | ☐ Yes ☐ No |  |
|  | Shall have a minimum of 3 sections: back section, seat section and leg section |  | ☐ Yes ☐ No |  |
|  | AC operation with backup battery (capacity for a minimum of 1 week of operation) |  | ☐ Yes ☐ No |  |
|  | Shall be equipped with electromechanically driven running gear and manual back up capability with a braking system that does not allow any movement when engaged. |  | ☐ Yes ☐ No |  |
|  | Operating tabletop with soft padded, easy to clean, flame and tear proof, anti static and chemical resistant covering. |  | ☐ Yes ☐ No |  |
|  | Shall include corded control unit with manual override. |  | ☐ Yes ☐ No |  |
|  | Motorized Height adjustment: 60 - 90 cm (Approximately) |  | ☐ Yes ☐ No |  |
|  | Motorized Trendelenburg / reverse Trendelenburg: ± 25° |  | ☐ Yes ☐ No |  |
|  | Motorized Lateral tilt: ± 20° |  | ☐ Yes ☐ No |  |
|  | Back section: +45° to –25° |  | ☐ Yes ☐ No |  |
|  | Leg section: ± 60° |  | ☐ Yes ☐ No |  |
|  | Load capacity (for the table and accessories): Static: 250 kg, Articulated: 180 kg |  | ☐ Yes ☐ No |  |
|  | Shall have a radio translucent top: |  | ☐ Yes ☐ No |  |
|  | Shall have adequate longitudinal access to facilitate C arm access |  | ☐ Yes ☐ No |  |
|  | Shall accept up to a 14” x 17” size x ray cassettes |  | ☐ Yes ☐ No |  |
|  | Accessories for general applications: |  | ☐ Yes ☐ No |  |
|  | Accessory rails (on both sides) |  | ☐ Yes ☐ No |  |
|  | Anesthesia screen |  | ☐ Yes ☐ No |  |
|  | Arm boards with mattress |  | ☐ Yes ☐ No |  |
|  | Body restraint and safety straps |  | ☐ Yes ☐ No |  |
|  | X-ray cassette tunnel |  | ☐ Yes ☐ No |  |
|  | Clamps for attachments |  | ☐ Yes ☐ No |  |
|  | The unit shall include all required accessories for full functionality |  | ☐ Yes ☐ No |  |
|  | Shall be CE marked and/or FDA approved |  | ☐ Yes ☐ No |  |
| 27 | Anesthesia Machine Minor with monitor |  |  |  |
|  | Name of Manufacturer | 1 | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable It is mandatory to send the brochures / data sheet of the product and identify the page number in those brochures to conform to the required specs. Fill all required details as mentioned above for all the cells below |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Valid ISO 13485 or 9001 or equivalent |  | ☐ Yes ☐ No |  |
|  | A microprocessor-controlled anesthesia machine with embedded electronic ventilator complete with all accessories and a patient monitor, for adult, pediatric, and Neonatal use |  | ☐ Yes ☐ No |  |
|  | Anesthesia Unit |  | ☐ Yes ☐ No |  |
|  | Closed circuit system with possibility to work in open circuit |  | ☐ Yes ☐ No |  |
|  | Gas delivery system: 3 gas delivery system (O2, N2O and air) with both inlets and connections and B.O.C probes for connection to central gas outlets (O2, N2O & Medical Air (4 bar),color coded and separate portable pin index cylinders for O2& N2O. |  | ☐ Yes ☐ No |  |
|  | Pressure gauges for input pressure indication for each gas |  | ☐ Yes ☐ No |  |
|  | Detachable patient module with CO2 absorbent canister |  | ☐ Yes ☐ No |  |
|  | Oxygen-flush range: not less than 25 L/min |  | ☐ Yes ☐ No |  |
|  | Breathing Module: Fully autoclavable module to deliver fresh gas to the patient (materials in contact with exhaled patient gases are autoclavable), it should be easily removable and reinstalled without tools |  | ☐ Yes ☐ No |  |
|  | Bellow module |  | ☐ Yes ☐ No |  |
|  | Safety systems |  | ☐ Yes ☐ No |  |
|  | Hypoxia guard (O2/N2O safety control): The unit should be equipped with integrated ratio system to maintain at least 25 % O2 in fresh gas whenever the oxygen supply pressure is reduced below normal or when accidental opening of only N2O flow with O2 valve closed. Also, it should be with alarm system |  | ☐ Yes ☐ No |  |
|  | oxygen failure or low pressure alarm system (visual and audible) with simultaneous stop of N2O gas delivery |  | ☐ Yes ☐ No |  |
|  | Adjustable up to at least 70 cm H2O Pressure Limiting valve (APL) |  | ☐ Yes ☐ No |  |
|  | Should have embedded electronic gas mixing |  | ☐ Yes ☐ No |  |
|  | Flow sensor: Reusable type |  | ☐ Yes ☐ No |  |
|  | The unit must include all necessary regulators and flow meters |  | ☐ Yes ☐ No |  |
|  | Inspiratory valve, expiratory valve & active gas scavenging port |  | ☐ Yes ☐ No |  |
|  | Suction regulator with all required hoses and probes (BOC) |  | ☐ Yes ☐ No |  |
|  | Ventilator |  | ☐ Yes ☐ No |  |
|  | Colored touchscreen digital display for ventilation parameters of at least 15 inches |  | ☐ Yes ☐ No |  |
|  | The machine and ventilator should be from the same manufacturer |  | ☐ Yes ☐ No |  |
|  | The equipment will be able to work with compressed air and oxygen pressures of about 3.5 – 5 bar |  | ☐ Yes ☐ No |  |
|  | Rechargeable internal battery (Minimum of 45 minutes operation when fully charged) |  | ☐ Yes ☐ No |  |
|  | Ventilation modes (at least the following): |  | ☐ Yes ☐ No |  |
|  | Manual / spontaneous |  | ☐ Yes ☐ No |  |
|  | Positive End Expiratory Pressure |  | ☐ Yes ☐ No |  |
|  | Pressure Controlled Ventilation |  | ☐ Yes ☐ No |  |
|  | Pressure Support Ventilation |  | ☐ Yes ☐ No |  |
|  | Volume Controlled Ventilation |  | ☐ Yes ☐ No |  |
|  | Synchronized Intermittent Mandatory Ventilation |  | ☐ Yes ☐ No |  |
|  | Programmable parameters (at least the following): |  | ☐ Yes ☐ No |  |
|  | Type of patient (adult, pediatric, neonate) |  | ☐ Yes ☐ No |  |
|  | Pressure values and limits |  | ☐ Yes ☐ No |  |
|  | Tidal volume 5-1400 ml or better |  | ☐ Yes ☐ No |  |
|  | Inspiration flow: 5 to 70 l/min |  | ☐ Yes ☐ No |  |
|  | Pressure limit: 15 to 70 cm H2O |  | ☐ Yes ☐ No |  |
|  | Maximum minute volume to be not less than 25 l/min |  | ☐ Yes ☐ No |  |
|  | Respiratory rate |  | ☐ Yes ☐ No |  |
|  | Oxygen concentration; FiO2 |  | ☐ Yes ☐ No |  |
|  | PEEP |  | ☐ Yes ☐ No |  |
|  | I: E Ratio 2: 1 to 4:1 or better |  | ☐ Yes ☐ No |  |
|  | Inspiratory pause: to include not less than the range 0 – 50 % |  | ☐ Yes ☐ No |  |
|  | Ventilation parameters monitoring (at least the following): |  | ☐ Yes ☐ No |  |
|  | Real time FiO2, capnograph, Anesthetic agents Isoflurane/ Sevoflurane & N2O (if not shown on the patient monitor) Tidal Volume, Minute Volume, Respiration Frequency, PEEP & Mean pressure |  | ☐ Yes ☐ No |  |
|  | MAC: It should have a display of MAC (Minimum Alveolar Concentration) |  | ☐ Yes ☐ No |  |
|  | Alarms |  | ☐ Yes ☐ No |  |
|  | Power Failure alarm Low O2 concentration Low/ High Tidal Volume O2 inlet supply failure |  | ☐ Yes ☐ No |  |
|  | Values and alarms of pressure data and settings |  | ☐ Yes ☐ No |  |
|  | Values and alarms of volume and flow data |  | ☐ Yes ☐ No |  |
|  | Trends and numerical data charts and curves |  | ☐ Yes ☐ No |  |
|  | Gas Module: |  | ☐ Yes ☐ No |  |
|  | There should be an anesthesia gas monitoring module that offers automatic anesthetic agent identification, etCO2 monitoring,FiO2,N2O etc.All necessary accessories to be included. |  | ☐ Yes ☐ No |  |
|  | Anesthesia Gas Scavenging System: |  | ☐ Yes ☐ No |  |
|  | Anesthetic gas scavenging system (Active) be present with standard connections. |  | ☐ Yes ☐ No |  |
|  | CO2 Absorber: |  | ☐ Yes ☐ No |  |
|  | A CO2 absorbent with a reusable canister should be present |  | ☐ Yes ☐ No |  |
|  | Vaporizer: |  | ☐ Yes ☐ No |  |
|  | Two independent system compatible vaporizers |  | ☐ Yes ☐ No |  |
|  | Selecta Tec type for fast and easy coupling or any equivalent type |  | ☐ Yes ☐ No |  |
|  | Offered vaporizers must include Isoflurane and Sevoflurane |  | ☐ Yes ☐ No |  |
|  | Should be easily mounted two vaporizers |  | ☐ Yes ☐ No |  |
|  | Test: The unit should be able to perform Leak Test,sensors calibration & start up Test |  | ☐ Yes ☐ No |  |
|  | vaporizers must be with key filling. Also, refilling adapters (keys) must be provided |  | ☐ Yes ☐ No |  |
|  | Accessories and Consumables for Anesthesia |  | ☐ Yes ☐ No |  |
|  | 3 m long gas hoses qty 3 (color coding green for oxygen, yellow for air, blue for nitrous protoxide) |  | ☐ Yes ☐ No |  |
|  | Sterilizable complete adult breathing circuits (with sampling luer-lock Y piece connector, filters, etc.) qty 2 |  | ☐ Yes ☐ No |  |
|  | Sterilizable complete pediatric breathing circuit (with sampling luer-lock Y piece connector, filters, etc.) qty 1 |  | ☐ Yes ☐ No |  |
|  | Sterilizable complete neonate breathing circuit (with sampling luer-lock Y piece connector, filters, etc.) qty 2 |  | ☐ Yes ☐ No |  |
|  | Set of 3 manual ventilation balloons |  | ☐ Yes ☐ No |  |
|  | 2x Jars of approximately 4 kg (or equivalent) of CO2 absorbent |  | ☐ Yes ☐ No |  |
|  | 10x disposable bacterial filters (adult & pediatric) |  | ☐ Yes ☐ No |  |
|  | Basic set of adult face masks, different sizes (small, medium, large) |  | ☐ Yes ☐ No |  |
|  | 2X Pediatric face mask |  | ☐ Yes ☐ No |  |
|  | 2 x Oxygen cell (FiO2) or 1 paramagnetic O2 sensor |  | ☐ Yes ☐ No |  |
|  | 2 x Flow sensor |  | ☐ Yes ☐ No |  |
|  | test lungs for adult, child and pediatrics |  | ☐ Yes ☐ No |  |
|  | Oxygen Gas Cylinders: (Qty. 2) (at least 5 litres) N2O Gas Cylinder (Qty. 1) (at least 5 litres) |  | ☐ Yes ☐ No |  |
|  | All standard accessories / consumables /parts required for the proper operation of the above system shall be included |  | ☐ Yes ☐ No |  |
|  | Trolley |  | ☐ Yes ☐ No |  |
|  | Easy to clean, resistant to corrosion, water, detergent soap, ethylic alcohol solution and to the hypochlorite of sodium |  | ☐ Yes ☐ No |  |
|  | Mounted on 4 castor antistatic wheels, at least two of them with brakes |  | ☐ Yes ☐ No |  |
|  | Drawers, shelf, work surface, absorber arms |  | ☐ Yes ☐ No |  |
|  | Back cylinders support |  | ☐ Yes ☐ No |  |
|  | Lateral rail for clamp connection |  | ☐ Yes ☐ No |  |
|  | Patient monitor holder |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | Compliant with IEC 60601 or equivalent |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |
|  | Anesthesia Patient Monitor |  | ☐ Yes ☐ No |  |
|  | Name of Manufacturer |  | ☐ Yes ☐ No |  |
|  | Model/ catalogue number |  | ☐ Yes ☐ No |  |
|  | Country of Origin for the offered model |  | ☐ Yes ☐ No |  |
|  | Equipment offered must be covered by at least a 2 years full warranty starting the date of installation and Final Acceptance |  | ☐ Yes ☐ No |  |
|  | Valid MDR (EU) 2017/745 or US FDA (510K) regulatory approval as Medical Devices or as “IVD” as applicable |  | ☐ Yes ☐ No |  |
|  | Modular |  | ☐ Yes ☐ No |  |
|  | ( ECG , NIBP, SpO2, RR, TEMP, IBP) Module w/ accessories,additional IBP module |  | ☐ Yes ☐ No |  |
|  | Options: |  | ☐ Yes ☐ No |  |
|  | Massimo technology option with accessories |  | ☐ Yes ☐ No |  |
|  | 2-channels Ibp option with accessories |  | ☐ Yes ☐ No |  |
|  | CO2 sidestream / main stream option with accessories |  | ☐ Yes ☐ No |  |
|  | Cardiac output PICO invasive option with accessories |  | ☐ Yes ☐ No |  |
|  | Continuous display on screen of patient ECG, respiration rate and heart rate, non-invasive blood pressure, body temperature and SpO2. |  | ☐ Yes ☐ No |  |
|  | Built in Battery |  | ☐ Yes ☐ No |  |
|  | Display: |  | ☐ Yes ☐ No |  |
|  | a) High resolution multi-color display (LCD) or TFT |  | ☐ Yes ☐ No |  |
|  | b) Size not less than 15" |  | ☐ Yes ☐ No |  |
|  | Simultaneously displayed information on the same screen: Numeric measured data and waveforms |  | ☐ Yes ☐ No |  |
|  | Visual and audible alarms with adjustable volume and limits |  | ☐ Yes ☐ No |  |
|  | PARAMETERS TO BE MONITORED: |  | ☐ Yes ☐ No |  |
|  | 1-ECG 10 LEAD, 5 LEAD, 3 LEAD with alarms |  | ☐ Yes ☐ No |  |
|  | 2-Respiration rate with alarms (Hi & low R.R. and apnea alarm) |  | ☐ Yes ☐ No |  |
|  | 3-NIBP: Systolic, Diastolic & mean – Mode: Automatic & Manual with alarms |  | ☐ Yes ☐ No |  |
|  | 4-Body temperature with alarms |  | ☐ Yes ☐ No |  |
|  | 5-SpO2 measurement with alarms |  | ☐ Yes ☐ No |  |
|  | 6-Heart Rate with alarms |  | ☐ Yes ☐ No |  |
|  | TRENDING Hrs.Not less than 72 hrs. for all monitored parameters |  | ☐ Yes ☐ No |  |
|  | Protections of all the functions against defibrillator discharges and electrosurgical units. |  | ☐ Yes ☐ No |  |
|  | Enclosure to have ingress protection level IPX1 or better. |  | ☐ Yes ☐ No |  |
|  | Accessories required for the above parameters |  | ☐ Yes ☐ No |  |
|  | it should be able to mount on the Anesthesia machine holder allowing the monitor to be manipulated |  | ☐ Yes ☐ No |  |
|  | Internal, easily accessible battery lasting for at least 2 hours |  | ☐ Yes ☐ No |  |
|  | Each unit must come complete with the following original accessories: |  | ☐ Yes ☐ No |  |
|  | ECG cable (3 leads), (with ECG trunk cable); Qty 2 |  | ☐ Yes ☐ No |  |
|  | ECG cable (5 leads), (with ECG trunk cable); Qty 2 |  | ☐ Yes ☐ No |  |
|  | ECG cable (10 leads), (with ECG trunk cable); Qty 2 |  | ☐ Yes ☐ No |  |
|  | Reusable Adult SpO2 sensor (with connector cable), 3 meters long; Qty 2 |  | ☐ Yes ☐ No |  |
|  | Reusable Pediatric SPO2 sensor (with connector cable), 3 meters long; Qty 2 |  | ☐ Yes ☐ No |  |
|  | Reusable Neonate SPO2 sensor (with connector cable), 3 meters long; Qty 2 |  | ☐ Yes ☐ No |  |
|  | Temperature probe; Qty 2 |  | ☐ Yes ☐ No |  |
|  | NIBP Neonate cuff with hose; Qty 2 |  | ☐ Yes ☐ No |  |
|  | NIBP Pediatric cuff with hose; Qty 2 |  | ☐ Yes ☐ No |  |
|  | NIBP Adult cuff with hose; Qty 2 |  | ☐ Yes ☐ No |  |
|  | Connectivity: HL7 License |  | ☐ Yes ☐ No |  |
|  | Electrical characteristics |  | ☐ Yes ☐ No |  |
|  | A 220-240V, 50Hz single-phase electrical source |  | ☐ Yes ☐ No |  |
|  | Compliant with IEC 60601 or equivalent |  | ☐ Yes ☐ No |  |
|  | Built-in protections against over-voltage and over-current line conditions. |  | ☐ Yes ☐ No |  |

Form C: Technical Quotation Form

RFQ reference no: [insert RFQ reference No.]

Name of Bidder: [insert name of Bidder]

Bidders are required to complete the Comparative Data Tables included in Section II: Schedule of Requirements to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.

Technical specifications for goods – Comparative Data Table

Delivery requirements –– Comparative Data Table

| UNOPS Requirements |  | Is quotation compliant? Bidder to complete | Details Bidder to complete |
| --- | --- | --- | --- |
| Delivery time | Bidder shall deliver the goods [120 calendar days] after Contract signature. | ☐ Yes ☐ No | Insert details |
| Delivery place and Incoterms rules | Delivery at Place DAP to Hospitals and Institutes mentioned in Bills of quantity tables- Form B, in case that the site is not ready, the supplier agrees to store the equipment in their warehouse until the site is prepared to receive the equipment. Incoterms rules as per Incoterms 2020. | ☐ Yes ☐ No | Insert details |
| Consignee details | This Shall include the Installation and Training requirements. Consignee will be the UNOPS and MoH representative | ☐ Yes ☐ No | Insert details |
| UNOPS Right to vary requirements | At the time the Contract is awarded, UNOPS reserves the right to vary the quantity of the goods and associated services specified above, provided this does not exceed +/- [20%] , without any change in the unit prices or other terms and conditions of the RFQ. | ☐ Yes ☐ No | Insert details |

00ax: +45 45 33 75 01

The offered goods and related services (if applicable) are in accordance with the required specifications and requirements specified in Section II: Schedule of Requirements.

☐ Yes   ☐ No

ANY DEVIATION MUST BE LISTED BELOW:

_________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

Name	: _____________________________________________________________

Title	: _____________________________________________________________

Date	: _____________________________________________________________

Signature	: _____________________________________________________________

If subcontracting is anticipated in the performance by the contractor, Procurement Official should insert the following:

List of subcontractors or suppliers

Bidder must identify the names of all subcontractors/suppliers who will be providing good/services under this contract and the type of work being subcontracted, if applicable.

- _[Full legal name and address of subcontractors]___________

- _________________________________________________

(C)_________________________________________________

I, the undersigned, certify that I am duly authorized by [insert full name of bidder] to sign this bid and bind [insert full name of bidder] should UNOPS accept this bid:

Name	: _____________________________________________________________

Title	: _____________________________________________________________

Date	: _____________________________________________________________

Signature	: _____________________________________________________________

# Form D: Previous Experience Form

RFQ reference no: [Insert UNOPS tender reference number]

Name of Bidder: [insert name of Bidder]

| Description of services/goods | Country | Total amount of Contract | Contract Identification and Title and Contact details of Client (Name, Address, telephone, email, fax) | Year project was undertaken |
| --- | --- | --- | --- | --- |

Name	: _____________________________________________________________

Title	: _____________________________________________________________

Date	: _____________________________________________________________

Signature	: _____________________________________________________________

# Form E: Self Disclosure Form

RFQ reference no: [insert RFQ reference No.]

Name of Bidder: [insert name of Bidder]

| Ref. | Critical area | Response | Comments |
| --- | --- | --- | --- |
| 1 | Has the entity or individual declared bankruptcy, or been involved in bankruptcy or receivership proceedings, or is there any judgment or pending legal action against them, which could impair operations in the foreseeable future; | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 2 | Has the entity or individual been found (or is awaiting a judgment or administrative decision) in breach of their obligations relating to the payment of taxes or social security contributions? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 3 | Has the entity or individual been found (or is awaiting a judgment or administrative decision) guilty of misconduct for the violation of applicable laws, regulations or ethical standards of the profession to which they belong? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 4 | Has the entity or individual engaged, or attempted to engage, in any Proscribed Practices in the past? For the purposes of this provision, “Proscribed Practices” are defined in the UNOPS Operational Instruction on Vendor Sanctions and include the following: - A corrupt practice: the offering, giving, receiving, or soliciting, directly or indirectly, anything of value to influence improperly the actions of another party; - A fraudulent practice: any act or omission, including a misrepresentation, that knowingly or recklessly misleads, or attempts to mislead, a party to obtain a financial or other benefit or to avoid an obligation; - A coercive practice: an act or omission that impairs or harms, or threatens to impair or harm, directly or indirectly, any party or the property of the party to improperly influence the actions of a party; - A collusive practice: an arrangement between two or more parties designed to achieve an improper purpose, including influencing improperly the actions of another party; - An unethical practice: Conduct or behaviour that is contrary to the conflict of interest, gifts and hospitality, post-employment provisions or other published requirements of doing business with UNOPS; and - Obstruction: Acts or omissions by a Vendor that prevent or hinder UNOPS from investigating instances of possible Proscribed Practices. | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 5 | Has the entity or individual unduly obtained, or attempted to unduly obtain, any confidential information in connection with this procurement process and any agreement that may be awarded as a result of this process? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 6 | Does the entity or individual have a conflict of interest, as provided in the Instructions to bidders, Article 3 “Bidder eligibility”, that may prevent them from entering into an agreement with UNOPS, or that may in any way jeopardize their performance in carrying out their obligations under the agreement? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 7 | Has the entity or individual been found (or is awaiting a judgment or administrative decision) to be involved in any of the following: - Fraudulent practice; - Corrupt practice; - Affiliation with a criminal organization; - Money laundering; - Terrorist financing; - Child labour; or - Human trafficking? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 8 | Has the entity or individual had significant performance issues in the past that led to the early termination of a legal commitment or the application of damages or other contractual penalties, or that were discovered following inspections, audits or investigations? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 9 | Has the entity or individual committed any financial irregularity, resulting from an act or omission or the inclusion of an unjustified item of expenditure outside a contractual engagement that had the effect of prejudicing their use of public funds? | Select | [If the response is ‘Yes’, provide an explanation here.] |
| 10 | Has the entity or individual been found (or is awaiting a judgment or administrative decision) to have created an entity in a different jurisdiction with the intent of circumventing fiscal, social or any other legal obligations in the jurisdiction of their registered office, central administration or principal place of business? | Select | [If the response is ‘Yes’, provide an explanation here.] |

I, the undersigned, hereby declare that the information provided is true and correct. I also understand that any willful dishonesty may result in the refusal of this submission. I understand that engagement in any of the above critical areas may, at the discretion of UNOPS, automatically result in the exclusion from this procurement process.

| Name: |  |
| --- | --- |
| Title |  |
| Date: | Sep 10, 2024 |
| Signature |  |

[Stamp this form with the official stamp of the bidder]

Form F: Joint Venture Partner Information Form

[The Bidder shall fill in this Form in accordance with the instructions indicated below].

RFQ reference no: [insert RFQ reference No.]

Name of Bidder: [insert name of bidder]

Date: [insert submission date]

To be completed and returned with your Bid if the Bid is submitted as a Joint Venture/Consortium/Association.

| JV / Consortium/ Association Information |  |
| --- | --- |
| Name | [complete] |
| Names of each partner and contact information (address, telephone numbers, fax numbers, e-mail address) | [complete] |
| Name of leading partner (with authority to bind the JV, Consortium, Association during the Bidding process and, in the event a Contract is awarded, during contract execution) | [complete] |
| Proposed proportion of responsibilities between partners (in %) with indication of the type of the goods/services to be delivered by each | [complete] |

Signatures of all partners of the JV:

We hereby confirm that if the contract is awarded, all parties of the Joint Venture/Consortium/Association shall be jointly and severally liable to UNOPS for the fulfillment of the provisions of the Contract.

Name of partner: ________________________		Name of partner: _________________________

Signature: ______________________________		Signature: _______________________________

Date: _______________________				Date: ________________________

Name of partner: ________________________		Name of partner: _________________________

Signature: ______________________________		Signature: _______________________________

Date: _______________________				Date: ________________________

Form G: Manufacturer’s Authorization Form

A letter issued by the manufacturer authorizing the applicant to participate in this particular RFQ must be submitted with the bid in the format provided in this Form.

To be eligible for delivery of goods, the bidder must be either the manufacturer of the offered goods or a sole representative of the manufacturer to the United Nations. Should offers for a particular make and model be received from more than one appointed representative, UNOPS reserves the right to select only one.

RFQ reference no: [Insert UNOPS tender reference number]

Name of Bidder: [insert name of bidder]

Date: [insert submission date]

To: [bidder to insert]

WHEREAS

We [insert complete name of manufacturer], who are official manufacturers of [insert type of goods manufactured], having factories at [insert full address of manufacturer’s factories], do hereby authorize [insert complete name of bidder] to submit a bid the purpose of which is to provide the following goods, manufactured by us [insert name and or brief description of the goods], and to subsequently negotiate and sign the contract.

We hereby extend our full guarantee and warranty in accordance with Clause 13 of the General Conditions for Goods, with respect to the goods offered by the above firm.

Signed: [insert signature(s) of authorized representative(s) of the manufacturer]

Name: [insert complete name(s) of authorized representative(s) of the manufacturer]

Title: [insert title]

Dated on ____________ day of __________________, _______ [insert date of signing]